DC997A Test Procedure

Required Equipment:

- Two RF signal generators HP8640 or HP8644B for encode clock and analog input.

- Power supply, 3.3V, 1Amp.

- DVMs, Two to measure Demo Board voltage and current.

- FastDAACS DC890B with 6V wallwart or external 6V @1.5Amp power supply.

- Computer configured with PScope Demo Board Test Version and EE_Programmer.exe for configuring Demo Board ID PROM.

- 250MHz, 210MHz or 170MHz encode clock source filters per table.

- 70MHz Signal filter.

Test procedure:

NOTE:  Always make sure the power supply is OFF before connecting it to the Demo board.  The charged output capacitor will defeat the protection of the current limit if the supply is ON.

Before beginning the test, configure the on board ID EEROM using EE_Programmer.exe.  This will allow PScope to auto-detect the board.  The board should be un-powered.  The DC890 provides power to the ID EEROM.

Adjust the power supply to 3.3V ± 100mV and set the current limit to 500mA.  Turn the power supply off

Verify all Jumpers are configured as follows

- SHDN		GND

- OE#		GND

- SENSE	+2.5V

- MODE		2/3 for LTC2242, for all other parts select VDD.

Connect the 3.3V power supply to the +3.3V and Gnd turrets.

Attach the Demo Board to the DC890B.

Verify that the DC890B is receiving external power from the wallwart.

Turn on the 3.3V power supply to the  DC997A demo board.  Measure the voltage on the 2.5V [option] test point.  This should read 2.5V ±50mV.

The Quiescent current with no clock or input signal should read:

- LTC2242CUP		350mA ± 50mA

- LTC2241CUP		290mA  ± 50mA

- LTC2240CUP		240mA  ±50mA

Connect the filtered Encode Clock with the signal generator set to +13dBm to the Encode SMA.  Use the appropriate filter selected from the chart below.

- LTC2242CUP		250Mhz

- LTC2241CUP		210MHz

- LTC2240CUP		170MHz

Connect the filtered 70 MHz signal generator to the AIN SMA

Select the correct part on PScope test, initiate data collection and adjust levels and frequencies as directed by the software or as shown in the table below.  The software will indicate Pass/Fail on completion of the test.

If using standard PScope, the pass fail criteria are in the table below – set PScope for 8192 pnt FFT and blkmm-harris 92dB windowing:

| Device | Clock Freq [+13 dBm] | Test Freq [-1 dBFS] | SNR dBc | SFDR dBc |
| --- | --- | --- | --- | --- |
| LTC2242-12 | 250MHz | 69.976807MHz | >64 | >70 |
| LTC2242-10 | 250MHz | 69.976807MHz | >60 | >70 |
| LTC2241-12 | 210MHz | 70.777587MHz | >64.5 | >70 |
| LTC2241-10 | 210MHz | 70.777587MHz | >60 | >70 |
| LTC2240-12 | 170MHz | 69.996338MHz | >65 | >70 |
| LTC2240-10 | 170MHz | 69.996338MHz | >60 | >70 |

Mark the Demo Board using an Ultra Fine Point Sharpe with the:

- device Type

- BIT Resolution

- Sample Rate (in Msps).

Package the demo board for delivery to samples.

QED.