## Ch
标准电阻首次校准 |  |  |  | SR104内部温度计算 |  |  | lymex, Feb 2016
标准电阻, Rs | 值 | 单位 | 备注 | 参数名称 | 参数值 | 单位 | 备注
标称值 | 10000 | Ohm |  | 标称值 | 10000 | Ohm | 温度传感器的标称值
校准日期 | 1984-08-21 |  | 对于SR104而言为NBS/NIST测试日期 | 参考温度下的偏差 | 0% |  | 这个数据标在出生证上
校准偏移量 | 0.54 | ppm |  | 参考温度，Tref | 23 | deg C
校准值 | 10000.0054 | Ohm | 如果首次校准数据丢失，删除“值”对应的数据 | 温度系数 | 0.1% | 每 deg C
校准温度，Tref | 23 | deg C | Tref, 参考温度 | 传感器测量值 | 9950 | Ohm
温度系数Alpha | 0.08 | ppm/C |  | 内部温度 | 18 | deg C
温度系数Beta | -0.028 | ppm/C2
标准电阻新近校准
校准日期 | 2015-12-18
校准偏移量 | 1.25 | ppm
校准值 | 10000.0125 | Ohm | 如果没有新近校准，删除“值”对应的数据
平均年漂移 | 0.022666 | ppm/year
测量数据
标准电阻, Rs | 值 | 单位 | 备注
测量日期 | 2016-03-21
测量偏偏移量 | 9.59 | ppm | 测量得到的与标称值之间的差异
测量值 | 10000.0959 | Ohm | 测量得到的与标称值之间的差异
测量温度 Tm1 | 18 | deg C | 标准电阻内部温度
未知电阻，Rx
测量偏偏移量 | 23 | ppm | 测量得到的与标称值之间的差异
测量值 | 10000.23 | Ohm | 测量得到的与标称值之间的差异
测量温度 Tm2 | 18 | deg C | 未知电阻内部温度
测量的温度系数Alpha |  | ppm/C | 此系数若为未知，删除“值”对应的数据
测量的温度系数Beta |  | ppm/C2 | 此系数若为未知，删除“值”对应的数据
结果
标准电阻, Rs | 值 | 单位 | 备注 | 备注：
在Tref温度下预测值 | 10000.0126 | Ohm | 标准电阻在参考温度下的阻值 | 绿色背景单元需要输入
在Tm1温度下的计算值 | 10000.0016 | Ohm | 标准电阻在测量温度Tm1下的阻值 | 黄色背景单元为计算结果，内有公式
未知电阻，Rx |  |  |  | 紫色背景，如果没有对应数据，对应的值应删除
在Tm2温度下的计算 | 10000.1357 | ppm | 未知电阻在测量温度Tm下的阻值
在Tref温度下的计算 | 10000.1357 | Ohm | 未知电阻在参考温度Tref下的阻值
Note on C6: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on C13: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on D15: ZhangLM: 此单元不要动
Note on C19: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on C23: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on C24: ZhangLM: This may be different than the standard

## En
Initial Calibration of the Standard Resistor |  |  |  | Temperature Calculation of SR104 |  |  | lymex, Feb 2016
The Standard, Rs | Value | Unit | Remarks | Parameter Name | Value | Unit | Remarks
Nominal Value | 10000 | Ohm |  | Nominal Value | 10000 | Ohm | of the temperature sensor
Calibrated date | 1982-03-12 |  | NBS/NIST test date for SR104 | Deviation at Tref | 0% |  | this is on the certificate
Calibrated deviation | 0.2 | ppm |  | Ref temp Tref | 23 | deg C
Calibrated Value | 10000.002 | Ohm | Delete the 'value' if no initial calibration | Tempco | 0.1% | per deg C
Calibrated Temp Tref | 23 | deg C | Tref, reference temperature | Sensor value | 9975 | Ohm | your measurement
Tempco Alpha | -0.1 | ppm/C |  | Temperature | 20.5 | deg C | result
Tempco Beta | -0.028 | ppm/C2
Recent Calibration of the Standard Resistor
Calibrated date | 2013-07-01
Calibrated deviation | 2.4 | ppm
Calibrated Value | 10000.024 | Ohm | Delete the 'value' if no recent calibration
Annual Drift | 0.070275 | ppm/year | Average Annual Drift
Measurement Data
The Standard, Rs | Value | Unit | Remarks
Measured date | 2016-03-10
Measured deviation | 12.3 | ppm | the instrument reading of your measurement
Measured value | 10000.123 | Ohm | the instrument reading of your measurement
Measured temp Tm1 | 20.5 | deg C | internal temperature of the standard
The Unknown Rx
Measured deviation | 15.5 | ppm | the instrument reading of your measurement
Measured value | 10000.155 | Ohm | the instrument reading of your measurement
Measured temp tm2 | 20.5 | deg C | internal temperature of the unknown
Tempco Alpha | 0.15 | ppm/C | Delete the 'value' if not known
Tempco Beta | -0.03 | ppm/C2 | Delete the 'value' if not known
The Result
The Standard, Rs | Value | Unit | Remarks | Note:
Predicted at Tref | 10000.0259 | Ohm | predicted value of your standard at Tref | Green cells are for input
Value at Tm1 | 10000.0266 | Ohm | the result of Rs at Tm | Yellow cells are the calculated results
The Unknown Rx |  |  |  | Delete the 'value' if not known
Value at Tm2 | 10000.0586 | ppm | the result of Rx at Tm
Value at Tref | 10000.0643 | Ohm | the result of Rx at Tref
Note on C19: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on C23: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on D15: ZhangLM: No change for this cell
Note on C13: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation
Note on C24: ZhangLM: This may be different than the standard
Note on C6: ZhangLM: If thers is no deviation value, this cell will be green and should be typed-in with value, otherwise it is calculated from the deviation