## Sheet1
Fluke 515A Calibrator Calibration Data
Serial Number: | 3970000 | [5 empty] | Calibration Certificate No:
Page: 1 of 2 | [6 empty] | Calibration Date:
DC µVolt O/P and Linearity | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper |  |  | Temperature
µV Dial Setting |  |  | Lower Limit |  | Measured |  | Upper Limit
990 |  |  | 988 | µV | 990.33 | µV | 992 | µV |  | 58.25% | [11 empty] | l | [8 empty] | Pass | 1. Using Fluke 80AK-A and Type K cable input 23.0°C
888 |  |  | 886 | µV | 887.72 | µV | 890 | µV |  | 43% | [8 empty] | l | [11 empty] | Pass | 2. Enter meter SETUP mode (Press yellow button then Backlight button), adjust temperature offset so
777 |  |  | 775 | µV | 776.82 | µV | 779 | µV |  | 45.5% | [9 empty] | l | [10 empty] | Pass | UUT displays 23.0°C using the Blue shift button and backlight button
666 |  |  | 664 | µV | 665.65 | µV | 668 | µV |  | 41.25% | [8 empty] | l | [11 empty] | Pass | along with the up/down arrow keys. Press yellow button followed by Backlight button to save offset.
555 |  |  | 553 | µV | 554.98 | µV | 557 | µV |  | 49.5% | [9 empty] | l | [10 empty] | Pass | 3. Proceed to measurement points
444 |  |  | 442 | µV | 444.07 | µV | 446 | µV |  | 51.75% | [10 empty] | l | [9 empty] | Pass
333 |  |  | 331 | µV | 333.84 | µV | 335 | µV |  | 71% | [14 empty] | l | [5 empty] | Pass | AC Tests, Voltage and Current
222 |  |  | 220 | µV | 222.91 | µV | 224 | µV |  | 72.75% | [14 empty] | l | [5 empty] | Pass
111 |  |  | 109 | µV | 112.36 | µV | 113 | µV |  | 84% | [16 empty] | l |  |  |  | Pass | The Range, Freq column refers to the UUT range (using manual ranging) and the Freq of the applied
0 |  |  | -3 | µV | 1.27 | µV | 3 | µV |  | 71.17% | [14 empty] | l | [5 empty] | Pass | test voltage or current
10 Volt DC Range and Linearity Test | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range | Multiplier sw |  | Lower Limit |  | Measured |  | Upper Limit
10 V | X | 1 | 9.9997 | V | 10 | V | 10.0003 | V |  | 49.17% | [9 empty] | l | [10 empty] | Pass
10 V | X | 0.9 | 8.9997 | V | 9 | V | 9.0003 | V |  | 50.56% | [10 empty] | l | [9 empty] | Pass
10 V | X | 0.8 | 7.9998 | V | 8 | V | 8.0002 | V |  | 48.96% | [9 empty] | l | [10 empty] | Fail
10 V | X | 0.7 | 6.9998 | V | 7 | V | 7.0002 | V |  | 48.81% | [9 empty] | l | [10 empty] | Fail
10 V | X | 0.6 | 5.9998 | V | 6 | V | 6.0002 | V |  | 46.94% | [9 empty] | l | [10 empty] | Fail | Frequency Test
10 V | X | 0.5 | 4.9999 | V | 5 | V | 5.0001 | V |  | 48.67% | [9 empty] | l | [10 empty] | Fail
10 V | X | 0.4 | 3.9999 | V | 4 | V | 4.0001 | V |  | 51.25% | [10 empty] | l | [9 empty] | Pass | 1. Leaving UUT on ACV, press Hz%ms button once for Hz
10 V | X | 0.3 | 2.9999 | V | 3 | V | 3.0001 | V |  | 50% | [10 empty] | l | [9 empty] | Pass | 2. Apply 1V 20 KHz
10 V | X | 0.2 | 1.9999 | V | 2 | V | 2.0001 | V |  | 54.17% | [10 empty] | l | [9 empty] | Pass
10 V | X | 0.1 | 0.99997 | V | 1 | V | 1 | V |  | 51.67% | [10 empty] | l | [9 empty] | Pass | Duty Cycle
AC Voltage Accuracy | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper |  |  | 1. Leaving UUT on ACV, press Hz%ms button once more for %
Range, Freq | Test Voltage |  | Lower Limit |  | Measured |  | Upper Limit | [24 empty] | 2. Apply 2 Vpp 500 Hz 30% duty cycle
1 V 400 Hz | 1 | V | 0.9995 | V | 0.99999 | V | 1.0005 | V |  | 49.5% | [9 empty] | l | [10 empty] | Pass
10 V 400 Hz | 10 | V | 9.996 | V | 9.9995 | V | 10.004 | V |  | 43.75% | [8 empty] | l | [11 empty] | Pass
100 V 400 Hz | 100 | V | 99.94 | V | 100.0002 | V | 100.06 | V |  | 50.17% | [10 empty] | l | [9 empty] | Pass
10 V 4 kHz | 10 | V | 9.99 | V | 9.9997 | V | 10.01 | V |  | 48.3% | [9 empty] | l | [10 empty] | Pass
10 V 50 kHz | 10 | V | 9.99 | V | 10.0005 | V | 10.01 | V |  | 52.3% | [10 empty] | l | [9 empty] | Pass
Frequency Test | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range, Freq | Test Freq |  | Lower Limit |  | Measured |  | Upper Limit
10 V 400 Hz | 400 | Hz | 396 | Hz | 400.0059 | Hz | 404 | Hz |  | 50.07% | [10 empty] | l | [9 empty] | Pass
10 V 4 kHz | 4 | kHz | 3960 | Hz | 4000.512 | Hz | 4040 | Hz |  | 50.64% | [10 empty] | l | [9 empty] | Pass
10 V 50 kHz | 50 | kHz | 47500 | Hz | 50005.93 | Hz | 52500 | Hz |  | 50.12% | [10 empty] | l | [9 empty] | Pass
Fluke515A Rev. 0 Mar. 2017
Fluke 515A Calibrator Calibration Data
Serial Number: | 3970000 | [5 empty] | Calibration Certificate No: | [10 empty] | 0
Page: 2 of 2 | [6 empty] | Calibration Date: | [10 empty] | 0
AC Distortion | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range | Test |  | Lower Limit |  | Measured |  | Upper Limit
1 V 400 Hz | 1 V | 400 Hz | ---- |  | 0.0052 | % | 0.03 | % |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | Pass
10 V 400 Hz | 10 V | 400 Hz | ---- |  | 0.0025 | % | 0.03 | % |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | Pass
10 V 4 kHz | 10 V | 4 kHz | ---- |  | 0.0024 | % | 0.03 | % |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | Pass
10 V 50 kHz | 10 V | 50 kHz | ---- |  | 0.0064 | % | 0.05 | % |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | Pass
Resistance Accuracy | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range | Resistance |  | Lower Limit |  | Measured |  | Upper Limit | [24 empty] | Press Blue shift button for AC+DC
0 Ω | 0 | Ω | -150 | mΩ | 26.7 | mΩ | 150 | mΩ |  | 58.9% | [11 empty] | l | [8 empty] | Pass
10 Ω | 10 | Ω | 9.994 | Ω | 10.0592 | Ω | 10.006 | Ω |  | 543% | [20 empty] | Fail
100 Ω | 100 | Ω | 99.94 | Ω | 100.016 | Ω | 100.06 | Ω |  | 63.33% | [12 empty] | l | [7 empty] | Pass
1 kΩ | 1 | kΩ | 999.85 | Ω | 1000.8411 | kΩ | 1000.15 | Ω |  | 330.4% | [20 empty] | Fail
10 kΩ | 10 | kΩ | 9998.5 | Ω | 9999.964 | kΩ | 10001.5 | Ω |  | 48.8% | [9 empty] | l | [10 empty] | Pass | Press Blue shift button for AC+DC
100 kΩ | 100 | kΩ | 99.985 | kΩ | 100.002 | kΩ | 100.015 | kΩ |  | 56.67% | [11 empty] | l | [8 empty] | Pass
1 MΩ | 1 | MΩ | 0.99985 | MΩ | 1.0001 | kΩ | 1.0002 | MΩ |  | 98% | [19 empty] | l | Pass
10 MΩ | 10 | MΩ | 9.9925 | MΩ | 10.0001 | MΩ | 10.0075 | MΩ |  | 50.68% | [10 empty] | l | [9 empty] | Pass
1. Short connected, wait a min. of 60 sec for meter to settle
Line Regulation with 1 kΩ load | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range | 102 VAC | 128 VAC | Lower Limit |  | Measured |  | Upper Limit
10 V 400 Hz | 9.9997 | 9.9997 | -100 | µV | 0 | µV | 100.003 | µV |  | 50% | [9 empty] | l | [10 empty] | Pass
Load Regulation load/no load | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range | Load | No Load | Lower Limit |  | Measured |  | Upper Limit
10 V 400 Hz | 9.9998 | 9.9998 | -0.0004 | µV | 5e-05 | µV | 0.0004 | µV |  | 56.25% | [11 empty] | l | [8 empty] | Pass | 1. Open connectors, no leads attached
10 V 4 kHz | 9.9996 | 9.9998 | -0.0004 | V | 0.00017 | V | 0.0004 | V |  | 71.25% | [14 empty] | l | [5 empty] | Pass
10 V 50 kHz | 10.0001 | 10.0007 | -0.0008 | V | 0.00057 | V | 0.0008 | V |  | 85.62% | [17 empty] | l |  |  | Pass
100 V 400 Hz | 99.9981 | 99.9995 | -0.0015 | V | 0.0014 | V | 0.0015 | V |  | 96.67% | [19 empty] | l | Pass
100 V | 100.0001 | 100.0001 | -0.0005 | V | 0 | V | 0.0005 | V |  | 50% | [10 empty] | l | [9 empty] | Pass | 1. Remove leads from UUT for Open test
DC Voltage Range Test | [10 empty] | lower |  | Limits in 5% increments | [15 empty] | upper
Range | Multiplier sw |  | Lower Limit |  | Measured |  | Upper Limit
100 V | N | /A | 99.997 | V | 100 | V | 100.003 | V |  | 50.33% | [10 empty] | l | [9 empty] | Pass
10 V | X | 1 | 9.9997 | V | 10 | V | 10.0003 | V |  | 48.17% | [9 empty] | l | [10 empty] | Pass
1 V | X | 1 | 0.99997 | V | 1 | V | 1 | V |  | 53.33% | [10 empty] | l | [9 empty] | Pass
| Metrologist/Technician | [5 empty] | Signature
Fluke515A Rev. 0 Mar. 2017