## Sheet1
[13 empty] | Measured | @ 100Hz |  | Replacement | @ 100Hz
Clip Qty | Item | Clip Ref | Clip Description | Clip Mfg Part No | Clip Mfg | Actual Marking | Spec | Selected Part | Selected Mfg | Supplier | Comments | Notes | Value | D | ESR | Value | D | ESR
1 | Fan | B1 | FAN-TBAX 12.4-CFH 12VDC | 3160-0492 | HP | PABST TYP 812L | 12.4CFM, 12V nominal actually runs at 15V | Papst 612NLE | PABST | RS |  | (*8)
1 | Fan Connector |  | CONNECTOR-HOUSING POST-TYPE | 50-57-9402 | Molex |  |  | (as specified) | MOLEX | Farnell
2 | Fan Connector |  | CONTACT-CONNECTOR U/W-POST-TYPE FEMALE CRIMP | 16-02-0088 | Molex |  |  | (as specified) | MOLEX | Farnell
1 | Mains Inlet & Filter | FRM1 | FILTER-LINE CEE-22-TERMS | 9135-0167 | HP | SCHAFFNER FN 323-3-05 | 3A | FN9233-3-06 | SCHAFFNER | Farnell |  | (*4)
2 |  | (not on CLIP) | Torx bolts for mains filter |  |  |  | M3 x 12mm torx bolt, plus nut, locking washer etc |  |  | eBay | Replace rivets on mains filter with torx bolts (+ M3 nuts, locking washers etc)
2 | 2200uF/35V | A4C7/8 | CAPACITOR-FXD 2200UF 35V AL·ELECT | 0180-4033 | HP | as spec, 85C | 18d, 36h, axial, black, 7.5mm pins | UPM1V222MHD | Nichicon | Mouser | 16d, 36l, 7.5mm pins |  | 1940 | 0.1 | 0.085 | 1990 | 0.06 | 0.05
1 | 15KuF/16V | A4C9-15 | CAPACITOR-FXD 15KUF 16V AL-ELECT | 0180-4034 | HP | as spec, 85C, Nippon Chemicon | 36d, 26h, radial, black, 10mm pins | UKW1C153MRD | Nichicon | Mouser | 20d, 40l, 10mm pins (max 48mm height) | (*7) | 11920 | 0.26 | 0.035 | 16500 | 0.37 | 0.04
1 | 22uF/50V | A4C17 | CAPACITOR-FXD 22UF+50-10% 50VDC AL | 513D226F063BB3D | SPRAGUE | as spec, 85C | 8d, 12h, orange, 3.5mm pins | PLX1H220MCL1TD | Nichicon | Mouser |  |  | 23.6 | 0.02 | 1.37
2 | 680uF/50V | A6C1/11 | CAPACITOR-FXO 680UF 50 V AL ELCTLT | 0180-4032 | HP | as spec, 105C | 16d, 23h, brown, 7.5mm pins | UPW1J681MHD3TN | Nichicon | Mouser | 18d, 20h, 7.5mm pins, these leaked in TiN's machine, (max ~23mm height) | (*7) | 690 | 0.06 | 0.14 | 635 | 0.05 | 0.13
1 | 8300uF/35V | A6C8 | CAPACITOR·FXD 8300UF+75·10% 35VOC AL | 53D832G035JP7 | SPRAGUE | as spec, 85C, Sprague | 67l,26d, 34mOhm ESR, 4110mA ripple, 4.3mA leakage | 53D832G035JP6 | Vishay Sprague | Avnet | Difficult to find | (*9) | 8030 | 0.16 | 0.032 | 8150 | 0.23 | 0.046
2 | 47uF/63V | A6C200/1 | CAPACITOR·FXD 47UF+-20% 63VOC AL | 0180-3828 | HP | as spec, 85C | 8d, 15h,3.5mm pins | UPM1J470MPD | Nichicon | Mouser |  |  | 45 | 0.05 | 1.72
1 | 330uF/25V | A6C202 | CAPACITOR·FXD 330UF+·20% 25VOC AL | 0180-3830 | HP | as spec, 85C | 10d,15h, 5mm pins | UPM1E331MPD6TD | Nichicon | Mouser |  |  | 340 | 0.12 | 0.57
1 | 220uF/16V | A6C203 | CAPACITOR-FXD 220UF+-20% 16VDC AL | 0180-3829 | HP | as spec, 85C | 8d,14h, 3.5mm pins | UPM1C221MPD | Nichicon | Mouser |  |  | 205 | 0.12 | 0.97
1 | DS1220Y | A5U132 | IC STATIC RANDOM-ACCESS MEMORY (DS1220Y-150) | 1818-4335 | HP | DS1220Y-150 | 2K x 8 SRAM | DS1220AD-150+ | Maxim Dallas | Mouser | 'AD' part instead of 'Y' part even though slightly different timings | (*2)(*5)
1 |  | (not on CLIP) | 24 pin DIP socket |  |  |  | 24 pin DIP socket | 2-1571552-8 | TE Connectivity / AMP | Mouser
2 | DS1230Y | A5U121/2 | IC DS1235Y-150 256K NONVOLATILE RAM | 1818-4336 | HP | DS1230Y-150+ (not original) | 32K x 8 SRAM | DS1230Y-150+ | Maxim Dallas | Mouser | DS parts were soldered in on early CLIP | (*1)(*3)(*6)
2 |  | (not on CLIP) | 28 pin DIP socket |  |  |  | 28 pin DIP socket | 828-AG11D-ESL-LF | TE Connectivity / AMP | Mouser | DS parts were soldered in on early CLIP
6 | EPROM | A5U110-15 | EPROM·PRGM27C512 REV4 | 03458-8884x | HP |  | TMS27C512 series 65536 by 8-bit UV EPROMs | TMS27C512-2JL | TI | eBay | Socketed even on early CLIP, difficult to find
Notes:
(*1) - December '89 CLIP listed data SRAM as DS1235Y-150, but my machine had replacement DS1230Y parts soldered in
(*2) - DS1220Y SRAM can be replaced by DS1220AD and ST part # M48Z12 (according to Maxim)
(*3) - DS1230Y SRAM can be replaced by ST part # M48Z35Y or TI part # BQ4011YMA (according to Maxim)
(*4) - Keysight actually shows a FN9222-3-06 replacement part, 373uA leakage, 0.1uF, 2.5mH
(*5) - Keysight actually shows a DS1220Y-120+
(*6) - Keysight actually shows DS1230Y both 120+ and 150+
(*7) - Height of replacements limited by the case dimensions
(*8) - Keysight actually shows a Pabst 622 L
(*9) - Avnet had stock in 2015