Memory Information and Procedure

Product:	34980A Multifunction Switch/Measure Mainframe and Modules, L44xxA Series LXI Switch and Control Instruments

Date:	January 28, 2013

Mainframe 34980A,

Mainframe L4421A/33A/37A/45A/50A/51A/52A,

Mainframe L4490A/91A:

| Memory Type: EEPROM | Memory Size: 2KB |
| --- | --- |
| Memory Function: I/O hardware addresses, boot parameters, LAN¦GPIB configuration settings |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): No |
| Memory Erase Processes: SYST:SEC:IMM clears all instrument memory (except instrument operation firmware (not user modifiable), I/O hardware addresses (not user modifiable), boot parameters (not user modifiable), and calibration parameters (refer to the “Calibration Security” section in the 34980A Service Guide manual for how to clear these)) and then cycles power to the instrument. |  |

| Memory Type: Flash | Memory Size: 8MB |
| --- | --- |
| Memory Function: Instrument states, channel labels, sequences, instrument operation firmware (not user modifiable) |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): No |
| Memory Erase Processes: SYST:SEC:IMM |  |

| Memory Type: SDRAM | Memory Size: 32MB |
| --- | --- |
| Memory Function: Measurement data buffer |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

| Memory Type: SRAM | Memory Size: 4KB |
| --- | --- |
| Memory Function: Boot-up information |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| - Memory Erase Processes: Cycle Power |  |

DMM:

| Memory Type: EEPROM | Memory Size: 4KB |
| --- | --- |
| Memory Function: Serial EEPROM containing DMM calibration constants and miscellaneous nonvolatile DMM settings. |  |
| User Modifiable (Y/N): Yes, mostly via calibration process | Volatile (Y/N): No |
| Memory Erase Processes: Can only be done by Agilent service. Renders DMM uncalibrated. |  |

| Memory Type: SRAM | Memory Size: 32KB |
| --- | --- |
| Memory Function: DMM operational data |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

| Memory Type: Flash | Memory Size: 256KB |
| --- | --- |
| Memory Function: DMM operation firmware |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: Firmware upgrade |  |

34921A, 34922A, 34923A, 34924A, 34925A, 34931A, 34932A, 34933A, 34937A, 34938A, 34939A, 34941A, 34942A, 34959A,

L4421A, L4433A, L4437A,

L4490A/L4491A Option 004

| Memory Type: EEPROM | Memory Size: 8KB |
| --- | --- |
| Memory Function: Relay counts, Initialization or configuration |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: Flash | Memory Size: 16KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 512B or 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

34934A

| Memory Type: EEPROM | Memory Size: 16KB |
| --- | --- |
| Memory Function: Relay counts, Initialization or configuration |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: Flash | Memory Size: 64KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

34945A, 34946A, 34947A, L4445A, L4490A/91A

| Memory Type: EEPROM | Memory Size: 8KB |
| --- | --- |
| Memory Function: Relay counts, Initialization or configuration |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: Flash | Memory Size: 64KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

34945EXT:

| Memory Type: EEPROM | Memory Size: 8KB |
| --- | --- |
| Memory Function: All channel/bank/tether configuration information |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): No |
| Memory Erase Processes: SYSTem:RMODule:RESet <slot> resets to factory defaults |  |

| Memory Type: Flash | Memory Size: 64KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

34950A and L4450A:

| Memory Type: EEPROM | Memory Size: 8KB |
| --- | --- |
| Memory Function: Initialization or configuration |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: Flash | Memory Size: 16KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

| Memory Type: Flash for FPGA | Memory Size: 1MB |
| --- | --- |
| Memory Function: Flash for fpga |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 256KB |
| --- | --- |
| Memory Function: waveform download/upload |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

34951A and L4451A:

| Memory Type: EEPROM | Memory Size: 8KB |
| --- | --- |
| Memory Function: Cal constants, cal string |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): No |
| Memory Erase Processes: Recalibrate to overwrite old data |  |

| Memory Type: Flash | Memory Size: 16KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 512B or 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

| Memory Type: Flash for FPGA | Memory Size: 128KB |
| --- | --- |
| Memory Function: Flash for fpga |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: DRAM | Memory Size: 2MB |
| --- | --- |
| Memory Function: Waveforms downloaded |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

34952A and L4452A:

| Memory Type: EEPROM | Memory Size: 8KB |
| --- | --- |
| Memory Function: Cal constants, cal string |  |
| User Modifiable (Y/N): Yes | Volatile (Y/N): No |
| Memory Erase Processes: Recalibrate to overwrite old data |  |

| Memory Type: Flash | Memory Size: 16KB |
| --- | --- |
| Memory Function: Flash Isp/iap for microprocessor program |  |
| User Modifiable (Y/N): No | Volatile (Y/N): No |
| Memory Erase Processes: (none) |  |

| Memory Type: SRAM | Memory Size: 512B or 1KB |
| --- | --- |
| Memory Function: Internal processor RAM not user accessible |  |
| User Modifiable (Y/N): No | Volatile (Y/N): Yes |
| Memory Erase Processes: Cycle Power |  |

Select one of the following:

Sanitization process is documented above

Sanitization process could be available with engineering resources

Sanitization process is not possible (i.e. sanitize the product by destruction