MODIFICATION KIT L220

# APPLICATION & DESCRIPTION

The following modification kit is for use with 2955 series radio test sets and 6500 Automatic Amplitude Analyser. Its use is required to enable the new type of Eldor EHT transformers  to be used as a service replacement on earlier build state CRT drive boards AC1 44828/890 (2955) and AT01/1 3964/137 (6500).

Its use is not required for the new AC1/1 boards 44830/033 fitted in 52955-500 from S/N 132705-001 and

52955-501 from S/N 132712-001 as these will accommodate the new Eldor transformer directly.

For use on 2955, see section 1 and section 2.

For use on 6500, see section 3.

The kit comprises of the following:

| Item | Description | M.I. Part Number |
| --- | --- | --- |
| 1) | 1 x L.O.P.T. + board | 44830/035 |
| 2) | 6 X SPRING CLIPS | 22328/863 |
| 3) | 1 X Mod & instructions (L220) | --------/----- |
| 4) | 1 X M3 NUT | 21882/108 |
| 5) | 1 X 3x8 mm SCREW | 21833/001 |
| 6) | 1 X 6.8nF CAPACITOR | 26582/496 |
| 7) | 1 X INSULATING PAD | 37491/450 |
| 8) | 1 X BC449 TRANSISTOR | 28457/821 |
| 9) | 100 mm DOUBLE SIDED FOAM TAPE | 17550/732 |
| 10) | 100 mm DOUBLE SIDED TAPE | 17550/424 |

PLEASE READ THE ENTIRE DOCUMENT BEFORE STARTING THE UPDATE PROCEDURE.

L220 MODIFICATION INSTRUCTIONS

There have been 2 major changes to the EHT transformer .

Most early units ( 2955 only and below s/n 132108-001) had Sarea transformers fitted (red). The company then started using Eldor transformers (Blue) who’s mechanical differences do not enable the screening can to be mounted in the same way. If you are using this kit to replace a Sarea transformer then carry out sections 1 and 2.

The Eldor transformer that has been used in later 2955’s and also  2955A/R/B  has now gone obsolete and the closest replacement available (still Eldor) has a number of  electrical differences. If you are using this kit to replace an earlier Eldor transformer then carry out section 2 only.

Section 1

Diagram 1

- Carefully remove and discard the Sarea L.O.P.T. but retain the screening can.

Note:	It is advisable to cut the legs of the old transformer to remove it. Cut the legs between the top 		of the PCB and the underside of the transformer.

The remainder of the pins can then be desoldered individually minimising the risk of damaging 		the delicate tracks on the PCB.

- Mark out holes  ‘A’ ‘B’ ‘C’ and ‘D’ ready for drilling as shown in diagram 3.

Note:-   It is advisable to centre punch the holes prior to drilling. In this way the board can be held up to the light to ensure that the holes will not interfere with any tracks on both the trackside and component side of the board. A useful check is to ensure that holes A & B fall on the vertical lines of the white board legend of the L.O.P.T. outline on the board. This can be seen when the board is held up against the light. The location of the holes may need to be altered slightly for some PM versions of the board.

- Once satisfied that the holes will be adequately clear of tracks they should be drilled through using a 2.2 mm dia. drill.

- With reference to diagram 2 press a spring clip through each of the holes. Ensure that the grooves of the clips run parallel with the vertical lines of the white board legend on the component side of the board. Carefully bend the tags over slightly on the trackside of the board. (The tags are made from hard material they can easily be broken and will not bend fully flat to the board.). To secure the tags firmly in position add ‘blobs’ of solder as shown in diagram 2.

Diagram 2

- On the trackside of the board add a sleeved wire link from the tag in hole ‘C’ to the grounded hole previously occupied by the old transformer. See diagram 3. This provides a ground for the screening can.

- Add the nut and screw supplied to the screening can as shown in diagram 1. This provides a threaded pillar for fixing the perspex safety cover when the board is reinserted into the instrument.

Once the modifications have been carried out in Section 2 the screening can is secured to the board by pressing the can into the securing spring clips.

NOTE:- Dimensions shown are from datum point ‘X’ which is the top left corner of the square surrounding hole 		‘y’ and not the edge of the circuit board.

Diagram 3  -  Part of AC1 board 44828/890 track-side

Section 2

The new Eldor transformer (M.I. Part 28231/417) has slightly different electrical and mechanical characteristics to the original Eldor transformer (M.I. Part 28231/402).

- When fitting the new Eldor transformer in place of the Sarea or earlier Eldor transformer the following modifications need to be carried out.

- With reference to the circuit ( diagram 4) replace TR2 2N2483 with new transistor BC449 (M.I. Part 28457/821).

- With reference to the circuit (diagram 4) and location diagram 5 or 5A add a new 6.8 nF capacitor C42 (M.I. Part 26582/496).  (On very early boards it may be necessary to locate C42 as in diagram 5A)

- With reference to diagram 6 position insulating pad (M.I. Part 37491/450) on PCB and attach to the PCB using two strips of double sided tape (M.I. Part 17550/424).

- Fit the new EHT transformer assembly complete with piggy back board to the CRT drive board PCB.

Note: See section 1 (step 1) detailing removal of the old transformer.

- Fitting the new transformer may affect the sensitivity of the front panel intensity control R33 should  be reselected if necessary.

TR2 TYPE CHANGED.

TR2 UNDERSIDE VIEW

C     B     E

Diagram 4 - part of circuit of 44828/890 board

Diagram 5 - track-side of 44828/890 board

Diagram 5A - track-side of 44828/890 (Early boards)

.

.

Diagram 6 - component-side of 44828/890 board

Section 3

Fitting the new Eldor EHT transformer to a 6500.

- The replacement transformer is supplied fitted to a piggy back board.  Make sure that the pins of the transformer do not protrude more than 2mm from the track side of the board.

- Solder the piggy back board into the position of the old transformer on the AT01/1 PCB.

- Form a wire link, insulated with PTFE sleeving, between pins 1 & 8 (as marked on the piggy back board) on the back of the main PCB.

- Mount the 6.8nF capacitor (item 6) in parallel with C24 (12nF).  Suggested method is as follows:

Remove C24 from the PCB.  Using the double sided tape, stick C24 and the new 6.8nF capacitor together, then stick both capacitors to the PCB such that they can be connected with tinned copper wire, as shown:

The remaining items in the mod kit are not used.

# Text boxes

ELDOR L.O.P.T (blue moulding)

SAREA L.O.P.T (red moulding)

NEW CAN

ORIGINAL CAN

Add solder ‘blobs’ as shown

X

Y

‘D’

‘B’

‘C’

‘A’

62mm

80mm

33mm

84mm

42mm

14 mm

14 mm

62mm

TR2 BC449

DOUBLE SIDED TAPE

12nF

68nF

TINNED COPPER WIRE

NEW CAPACITOR C42

TO BE LAID ON SIDE AND

SECURED TO PCB USING

DOUBLE SIDED TAPE (17550/732)

IN POSITION SHOWN.

New Capacitor body must

not extend above this line

NEW CAPACITOR C42

TO BE LAID ON SIDE AND

SECURED TO PCB USING

DOUBLE SIDED TAPE (17550/732)

IN POSITION SHOWN

WIRE LINK

WIRE LINK

WIRE LINK

( SECTION 1 DIAGRAM 2)

SPRING CLIP FOR CAN MOUNTING

ENSURE ALL WIRE LINKS

ARE ADEQUATELY

SLEEVED

D4

New Capacitor body must

not extend above this line

POSITION INSULATING PAD (37491/450)

ON PCB AS SHOWN AND ATTACH TO

PCB WITH TWO STRIPS OF DOUBLE SIDED

TAPE (17550/424)

ADD WIRE LINK WHEN TAG HAS BEEN FITTED IN HOLE ‘C’

THE DOTTED LINE  REPRESENTS  THE WHITE LEGEND OUTLINE OF THE L.O.P.T. THIS CAN BE SEEN IF BOARD HELD UP TO LIGHT

C42 ADDED