## voltage calibration
|  | step | commands | settings | magnitude | unit | instructions
|  | 0 |  |  |  |  | figure 2-3 from the service manual was interpreted for this calibration, where the Keithley 6500 DMM was used as connected in parallel
autorange | configure multimeter | 1 |  | filter: | 32 |  | set the DMM to DCV mode and autorange
|  |  |  | NPLC: | 5 | Line Cycles
|  | 2 |  | relative: | -2.3865e-07 | DCV | with the short installed, set relative measurements
Keithley 238 | configure calibration | 3 | N1X |  |  |  | set the Keithley 238 to source voltage and place it in operate mode
|  | 4 | C0X |  |  |  | send the following command to enter the calibration mode (note: current measure will not be displayed)
|  | 5 | C1X | wait: | 30 | s | with the short still installed, send the following command over the bus to perform common mode calibration
1.5 DCV range | +0 DCV offset | 6 | C2X |  |  |  | remove the short and send the following commands while recording the voltages from the DMM
|  | 7 |  | DMM: | 0.0001596 | DCV
|  | 8 | C3,1.5960E-4X
| +1.4 DCV gain | 9 | C4X
|  | 10 |  | DMM: | 1.4 | DCV
|  | 11 | C5,1.39995E+0X
| -0 DCV offset | 12 | C6X
|  | 13 |  | DMM: | -6.87e-05 | DCV
|  | 14 | C7,-6.87000E-5X
| -1.4 DCV gain | 15 | C8X
|  | 16 |  | DMM: | -1.4 | DCV
|  | 17 | C9,-1.40004E+0X
15 DCV range | +0 DCV offset | 18 | C10X
|  | 19 |  | DMM: | 0.00059 | DCV
|  | 20 | C11,5.90000E-4X
| +14 DCV gain | 21 | C12X
|  | 22 |  | DMM: | 14.0001 | DCV
|  | 23 | C13,1.40002E+1X
| -0 DCV offset | 24 | C14X
|  | 25 |  | DMM: | -0.0006 | DCV
|  | 26 | C15,-6.00000E-4X
| -14 DCV gain | 27 | C16X
|  | 28 |  | DMM: | -14.0014 | DCV
|  | 29 | C17,-1.40014E+1X
110 DCV range | +0 V offset | 30 | C18X
|  | 31 |  | DMM: | -0.00443 | DCV
|  | 32 | C19,-4.43000E-3X
| +100 DCV gain | 33 | C20X
|  | 34 |  | DMM: | 99.9745
|  | 35 | C21,9.99745E+1X
| -0 DCV offset | 36 | C22X
|  | 37 |  | DMM: | 0.000668
|  | 38 | C23,6.68000E-4X
| -100 DCV gain | 39 | C24X
|  | 40 |  | DMM: | -99.9728
|  | 41 | C25,-9.99728E+1X
Keithley 238 | complete calibration | 42 | C59X |  |  |  | check that the CAL LOCK switch is in the desired position for perminant storage
for perminant calibration: inward position
send the following command over the bus to exit calibration mode

## 1 A calibration
|  | step | commands | settings | magnitude | unit | instructions
|  | 0 |  |  |  |  | figure 2-4 from the service manual was interpreted for this calibration, where the Keithley 6500 DMM and the Keithley 238 SUM were used as connected in parallel
autorange | configure multimeter | 1 |  | filter: | 32 |  | set the DMM to DCI mode and set it to the 2 mA range or auto
|  |  |  | NPLC: | 5 | Line Cycles | alternative, directly connect a high-precision DMM if a calibrator source is unavailable
Keithley 238 | configure calibration | 2 | N1X |  |  |  | set the Keithley 238 to source current and place it in operate mode
|  | 3 | C0X |  |  |  | send the following command to enter the calibration mode (note: voltage measure will not be displayed)
1 A range | +0 A offset | 4 | C26X |  |  |  | send the following command to source +0 mA on the 1 A range
|  | 5 |  |  |  |  | set the SMU/calibrator to source 0 V and place the measure range to 100 uA
|  | 6 |  | DMM: | -7.3e-05 | DCI | from the display of the DMM, record the offet current
|  | 7 | C27,-7.3000E-5X |  |  |  | send the calculation current reading over the bus with the following command to calibrate the offset for the +1 A source and measurement range
| +0.9 A gain | 8 | C28X |  |  |  | send the following command to source +0.9 A
|  | 9 |  |  |  |  | set the SMU/calibrator to output -0.9 A on the 1 A range
|  | 10 |  | DMM: |  | DCI | from the display of the DMM, record the current reading
|  | 11 |  | DMM: | 0.9011 | DCI | calculate the actual current (I + 0.9 A) using the following calculation
|  | 12 | C29,9.0110E-1X |  |  |  | send the calculation current reading over the bus with the following command to calibrate the gain for the +1 A source and measurement range
| -0 A offset | 13 | C30X |  |  |  | send the following command to source -0 A on the 1 A range
|  | 14 |  |  |  |  | set the SMU/calibrator to source 0 V and place the measure range to 100 uA
|  | 15 |  | DMM: | -0.000214 | DCI | from the display of the DMM, record the offet current
|  | 16 | C31,-2.1400E-4X |  |  |  | send the calculation current reading over the bus with the following command to calibrate the offset for the -1 A source and measurement range
| -0.9 A gain | 17 | C32X |  |  |  | send the following command to source -0.9 A
|  | 18 |  |  |  |  | set the SMU/calibrator to output +0.9 A on the 1 A range
|  | 19 |  | DMM: |  | DCI | from the display of the DMM, record the current reading
|  | 20 |  |  | -0.90039 | DCI | calculate the actual current (I - 0.9 A) using the following calculation
|  | 21 | C33,-9.0040E-1X |  |  |  | send the calculation current reading over the bus with the following command to calibrate the gain for the -1 A source and measurement range
Keithley 238 | complete calibration | 22 | C59X |  |  |  | check that the CAL LOCK switch is in the desired position for perminant storage
for perminant calibration: inward position
send the following command over the bus to exit calibration mode

## 100 mA calibration
|  | step | commands | settings | magnitude | unit | instructions
|  | 0 |  |  |  |  | figure 2-4 from the service manual was interpreted for this calibration, where the Keithley 6500 DMM and the Keithley 238 SUM were used as connected in parallel
autorange | configure multimeter | 1 |  | filter: | 32 |  | set the DMM to DCI mode and set it to the 200 uA range or auto
|  |  |  | NPLC: | 5 | Line Cycles | alternative, directly connect a high-precision DMM if a calibrator source is unavailable
Keithley 238 | configure calibration | 2 | N1X |  |  |  | set the Keithley 238 to source current and place it in operate mode
|  | 3 | C0X |  |  |  | send the following command to enter the calibration mode (note: voltage measure will not be displayed)
100 mA range | +0 A offset | 4 | C34X |  |  |  | send the following command to source +0 mA on the 100 mA range
|  | 5 |  |  |  |  | set the SMU/calibrator to source 0 V and place the measure range to 100 uA
|  | 6 |  | DMM: | 4.79e-08 | DCI | from the display of the DMM, record the offet current
|  | 7 | C35,4.7900E-8X |  |  |  | send the current reading over the bus with the C35 command as follows to calibrate the offset for the 100 mA range
| +90 mA gain | 8 | C36X |  |  |  | send the following command to source +90 mA
|  | 9 |  |  |  |  | set the SMU/calibrator to output -90 mA on the 100 mA range
[6 empty] | DCI | from the display of the DMM, record the current reading
|  | 10 |  | DMM/SMU: | 0.08987 | DCI | calculate the actual current (I + 90 mA) using the following calculation
|  | 11 | C37,8.9870E-2X |  |  |  | send the calculation current reading over the bus with the following command to calibrate the gain for the +100 mA source and measurement range
| -0 A offset | 12 | C38X |  |  |  | send the following command to source -0 mA on the 100 mA range
|  | 13 |  |  |  |  | set the SMU/calibrator to source 0 V and place the measure range to 100 uA
|  | 14 |  | DMM: | 4.709e-08 | DCI | from the display of the DMM, record the offet current
|  | 15 | C39,4.7090E-8X |  |  |  | send the current reading over the bus with the C39 command as follows to calibrate the offset for the -100 mA range
| -90 mA gain | 16 | C40X |  |  |  | send the following command to source -90 mA
|  | 17 |  |  |  |  | set the SMU/calibrator to output +90 mA on the 100 mA range
|  | 18 |  |  |  | DCI | from the display of the DMM, record the current reading
|  | 19 |  | DMM/SMU: | -0.089839 | DCI | calculate the actual current (I - 90 mA) using the following calculation
|  | 20 | C41,-8.9839E-2X |  |  |  | send the calculation current reading over the bus with the following command to calibrate the gain for the -100 mA source and measurement range
Keithley 238 | complete calibration | 21 | C59X |  |  |  | check that the CAL LOCK switch is in the desired position for perminant storage
for perminant calibration: inward position
send the following command over the bus to exit calibration mode

## 1 nA to 10 mA calibration
|  | step | commands | settings | magnitude | unit | instructions
|  | 0 |  |  |  |  | figure 2-5 from the service manual was interpreted for this calibration, where the Keithley 6500 DMM and the Keithley 238 SUM were used as connected in parallel
autorange | configure multimeter | 1 |  | filter: | 32 |  | set the SMU/calibrator to standby at the 0 nA range
|  |  |  | NPLC: | 1 | Line Cycles | alternatively, measure a stable current source with an electrometer and then use this source to calibrate the gain for the instrument being calibrated
autorange | configure electrometer |  |  | filter: | 5
|  |  |  | NPLC: | 10
Keithley 238 | configure calibration | 2 | N1X |  |  |  | set the Keithley 238 to source current and place it in operate mode
|  | 3 | C0X |  |  |  | send the following command to enter the calibration mode (note: voltage measure will not be displayed)
nA range | 0.0000 nA offset | 4 | C42X |  | 20 | seconds | send the following command over the bus to calibrate the offset for the 1 nA measurement range while the SMU/calibrator is operating at 0 nA
| 0.9 nA gain | 5 |  |  | 9.003e-10 | DCI | connect the SMU/calibrator to an electrometer, set the source to 0.9 nA, and measure the value
|  | 6 | C43,9.0030E-10X |  | 20 | seconds | send the following command over the bus to calibrate the gain for the 1 nA range
|  |  |  | 0% | 9.003e-10 | DCI | varify the sourced current again with the electrometer
| 00.000 nA offset | 7 |  |  |  |  | place the SMU/calibrator on standby and set it to 9 nA
|  | 8 | C44X |  | 10 | seconds | send the following command over the bus to calibrate the offset for the 10 nA measurement range while the SMU/calibrator is operating at 0 nA
| 09 nA gain | 9 |  |  | 8.9994e-09 | DCI | connect the SMU/calibrator to an electrometer, set the source to 9 nA, and measure the value
|  | 10 | C45,8.9994E-9X |  | 10 | seconds | send the following command over the bus to calibrate the gain for the 10 nA range
|  |  |  | 0.03334% | 8.9994e-09 | DCI | varify the sourced current again with the electrometer
| 000.00 nA offset | 11 |  |  |  |  | place the SMU/calibrator on standby and set it to 90 nA
|  | 12 | C46X |  | 4 | seconds | send the following command over the bus to calibrate the offset for the 100 nA measurement range while the SMU/calibrator is operating at 0 nA
| 90 nA gain | 13 |  |  | 9.0064e-08 | DCI | connect the SMU/calibrator to an electrometer, set the source to 90 nA, and measure the value
|  | 14 | C47,9.0064E-8X |  | 4 | seconds | send the following command over the bus to calibrate the gain for the 100 nA range
|  |  |  | 0.0111% | 9.0064e-08 | DCI | varify the sourced current again with the electrometer
uA range | 0.0000 uA offset | 15 |  |  |  |  | place the SMU/calibrator in standby and set it to source 0 nA
|  | 16 | C48X |  |  |  | send the following command over the bus to calibrate the offset for the 1 uA measurement range while the SMU/calibrator is operating at 0 nA
| 0.9 uA gain | 17 |  |  | 9.0092e-07 | DCI | connect the SMU/calibrator to an electrometer, set the source to 0.9 uA, and measure the value
|  | 18 | C49,9.0092E-7X |  |  |  | send the following command over the bus to calibrate the gain for the 1 uA range
|  |  |  | 0.0111% | 9.0092e-07 | DCI | varify the sourced current again with the electrometer
| 00.000 uA offset | 19 |  |  |  |  | place the SMU/calibrator on standby and set it to 9 uA
|  | 20 | C50X |  |  |  | send the following command over the bus to calibrate the offset for the 10 uA measurement range while the SMU/calibrator is operating at 0 nA
| 09 uA gain | 21 |  |  | 8.995e-06 | DCI | connect the SMU/calibrator to an electrometer, set the source to 9 uA, and measure the value
|  | 22 | C51,8.9950E-6X |  |  |  | send the following command over the bus to calibrate the gain for the 10 uA range
|  |  |  | 0.05559% | 8.995e-06 | DCI | varify the sourced current again with the electrometer
| 000.00 uA offset | 23 |  |  |  |  | place the SMU/calibrator on standby and set it to 90 uA
|  | 24 | C52X |  |  |  | send the following command over the bus to calibrate the offset for the 100 uA measurement range while the SMU/calibrator is operating at 0 nA
| 90 uA gain | 25 |  |  | 9.0046e-05 | DCI | connect the SMU/calibrator to an electrometer, set the source to 90 uA, and measure the value
|  | 26 | C53,9.0046E-5X |  |  |  | send the following command over the bus to calibrate the gain for the 100 uA range
|  |  |  | -0.01111% | 9.0046e-05 | DCI | varify the sourced current again with the electrometer
mA range | 0.0000 mA offset | 27 |  |  |  |  | place the SMU/calibrator in standby and set it to source 0 mA
|  | 28 | C54X |  |  |  | send the following command over the bus to calibrate the offset for the 1 mA measurement range while the SMU/calibrator is operating at 0 nA
| 0.9 mA gain | 29 |  |  | 0.00090072 | DCI | connect the SMU/calibrator to an electrometer, set the source to 0.9 mA, and measure the value
|  | 30 | C55,9.0072E-4X |  |  |  | send the following command over the bus to calibrate the gain for the 1 mA range
|  |  |  | -0.0111% | 0.00090072 | DCI | varify the sourced current again with the electrometer
| 00.000 mA offset | 31 |  |  |  |  | place the SMU/calibrator on standby and set it to 9 mA
|  | 32 | C56X |  |  |  | send the following command over the bus to calibrate the offset for the 10 mA measurement range while the SMU/calibrator is operating at 0 mA
| 09 mA gain | 33 |  |  | 0.0090045 | DCI | connect the SMU/calibrator to an electrometer, set the source to 9 mA, and measure the value
|  | 34 | C57,9.0045E-3X |  |  |  | send the following command over the bus to calibrate the gain for the 10 mA range
|  |  |  | 0% | 0.0090045 | DCI | varify the sourced current again with the electrometer
|  | 35 |  |  |  |  | pleace the SMU/calibrator in standby and set it to 0 mV
|  | 36 |  |  |  |  | place the SMU/calibrator in operate
|  | 37 | C58X |  | 4 | minutes | send the following command over the bus to calibrate the 1 nA through 10 mA source ranges
Keithley 238 | complete calibration | 21 | C59X |  |  |  | check that the CAL LOCK switch is in the desired position for perminant storage
for perminant calibration: inward position
send the following command over the bus to exit calibration mode