##### Series 2600 System SourceMeter® Instruments

v1.3.3 Firmware Release Notes

Keithley Instruments, Inc.

28775 Aurora Road

Cleveland, Ohio 44139-1891

440-248-0400  Fax: 440-248-6168

http://www.keithley.com

# Contents

Contents	1

General Information	3

Supported Models	3

Installation Instructions	3

Upgrade Considerations for Models 2601/2602	3

Upgrade Considerations for Models 2611/2612	4

Upgrade Considerations for Models 2635/2636	4

v1.3.3 Release	5

Overview	5

Compatibility Concerns	5

Critical Fixes	5

Enhancements	5

Non-critical Fixes	5

Known Issues	6

v1.3.0 Release	12

Overview	12

Compatibility Concerns	12

Critical Fixes	12

Enhancements	13

Non-critical Fixes	16

v1.2.0 Release	22

Overview	22

Critical Fixes	22

Enhancements	22

Non-critical Fixes	24

v1.1.4 Release	27

Overview	27

Critical Fixes	27

Enhancements	27

Non-critical Fixes	27

v1.1.3 Release	28

Overview	28

Critical Fixes	28

Non-critical Fixes	28

v1.1.2 Release	29

Overview	29

v1.1.1 Release	30

Overview	30

Critical Fixes	30

Enhancements	30

Non-critical Fixes	31

v1.1.0 Release	33

Overview	33

v1.0.X Releases	33

Overview	33

# General Information

## Supported Models

This firmware is intended for use on the following Keithley products:

2601, 2602,

2611, 2612,

2635, 2636

## Installation Instructions

Firmware installation instructions are given in the “Flash firmware upgrade” section in Chapter 13 of the Series 2600 System SourceMeter Instruments Reference Manual (Document Number: 260x-901-01).  This manual is available online at http://www.keithley.com/support/data?asset=50299.  Alternatively, Keithley factory upgrades can be arranged by calling your local Keithley support office.

## Upgrade Considerations for Models 2601/2602

The following table outlines the considerations that should be made when deciding to upgrade a Model 2601/2602 to v1.3.3 or not from any previous version.

| Consideration | From v1.0.2 | From v1.0.3 v1.0.4 | From v1.0.5 v1.1.1 v1.1.2 v1.1.3 v1.1.4 | From v1.1.5 | From v1.1.6 |
| --- | --- | --- | --- | --- | --- |
| Recalibration Required | No | No1 | No1 | No1 | No1 |
| Re-qualification Suggested | Yes | No | No | No | No |
| Backward Compatibility Concerns | Yes | YesError: Reference source not found | YesError: Reference source not found | YesError: Reference source not found | YesError: Reference source not found |
| Should you upgrade? | Review | Yes | ReviewError: Reference source not found | Yes | ReviewError: Reference source not found |

## Upgrade Considerations for Models 2611/2612

The following table outlines the considerations that should be made when deciding to upgrade a Model 2611/2612 to v1.3.3 or not from any previous version.

| Consideration | From v1.2.0 |
| --- | --- |
| Recalibration Required | No1 |
| Backward Compatibility Concerns | YesError: Reference source not found |
| Should you upgrade? | Yes |

## Upgrade Considerations for Models 2635/2636

The following table outlines the considerations that should be made when deciding to upgrade a Model 2635/2636 to v1.3.3 or not from any previous version.

| Consideration | From v1.3.0 |
| --- | --- |
| Recalibration Required | No |
| Backward Compatibility Concerns | No |
| Should you upgrade? | ReviewError: Reference source not found |

# v1.3.3 Release

## Overview

Version 1.3.3 is a maintenance release of the Series 2600 firmware.  This firmware includes some non-critical fixes.  Note that versions 1.3.1 and 1.3.2 of the Series 2600 firmware were never released.

This version of firmware introduces support for the KIPulse script on the Model 2601 and Model 2602.  The KIPulse script was originally supported in v1.2.0 for the model 2611 and model 2612 but v1.2.0 was never approved for use on the model 2601 or model 2602.  See the updated Series 2600 manual for details on the KIPulse factory script.

## Compatibility Concerns

Version 1.3.0 firmware introduced changes that may affect compatibility with automated calibration software used to calibrate a Series 2600 instrument.  This firmware includes these changes and may be of concern when upgrading from a version of firmware earlier than v1.3.0.  See enhancement PR31753 in the Enhancements portion of the v1.3.0 release section for more details.

## Critical Fixes

There were no critical fixes included in this release.  See the Non-critical Fixes section for more information about release content.

## Enhancements

There were no enhancements included in this release.  See the Non-critical Fixes section for more information about release content.

## Non-critical Fixes

Models Affected:

PR34445

PR34446

PR34447

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

An SMU will return erroneous overflow indications making filtered measurements under the following conditions:  A measurement is made on one function, I or V.  Then the measure range for that function is changed such that the measurement just made would overflow the new range.  Then a measurement is made on the complementary function.

Resolution:

This issue has been corrected.

Models Affected:

PR34480

2611, 2612, 2635, 2636

Symptom:

The KIPulse script will generate an error about operating outside the Standard Operating Area (SOA) even though the pulse is within the SOA if the SMU is configured to source more than 20V and it is configured to source a low voltage pulse with a current limit greater than 1A.  The unit may then hang when attempting to turn the output off.

Resolution:

This issue has been corrected.

Models Affected:

PR34693

PR34728

PR34729

2611, 2612, 2635, 2636

Symptom:

When using the KIPulse script, if the specified pulse level or the compliance level exceeds a range boundary by 1% or less, the script will not choose the appropriate range.  Furthermore, if the levels specified are outside of the Standard Operating Area (SOA) of the instrument, the script may not detect the outside of SOA condition.  If the script does not detect the outside SOA condition the script will likely not output a pulse at the correct output level and may also generate various error messages.

Resolution:

This issue has been corrected.

## Known Issues

Models Affected:

PR26312

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The digio trigger pulse width is a minimum guarantee.  Due to software overhead, trigger pulse widths are typically approximately 70uS longer than requested in the ICL commands.  This is significant overhead for pulses that only need to be 10uS or so.

Models Affected:

PR26390

PR26444

PR26451

PR33257

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The unit may fail to operate correctly after an out of memory error.  The unit may ignore commands sent over the command interface(s) as well as ignore front panel operations.  To avoid out of memory issues, the user should not exceed 3MB of memory consumption for the TSP® runtime environment and reading buffers combined.  The manual outlines how to determine the amount of memory consumed for these.  The meminfo() function (see enhancement PR27752 in the v1.1.1 enhancements section below for more information on meminfo) can also be used to monitor the actual free memory remaining.  When the free memory drops below 500 kilobytes, there is a danger of the unit encountering an out of memory error.

Models Affected:

PR26758

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Currently there is no way for the user to view the contents of the front panel LOAD menu.  This is especially problematic for nonvolatile menu items as they are recreated after a power cycle.  Once a LOAD menu item has been created, the only portion of it a user can see is what appears on the front panel.  It is inconvenient to make a list of all items by visually inspecting the LOAD menu contents.  It is also impossible to see which items are stored in nonvolatile memory and it is impossible to see what action each menu item performs.

Models Affected:

PR26845

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The instrument does not properly detect Query Unterminated errors.  If the instrument is addressed to talk and there are no pending commands that might generate output, the instrument will fail to detect and generate an error in this situation.  This situation can be detected by the user when a GPIB read operation times out.

Models Affected:

PR26846

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Keithley-defined global variables (logical instruments, etc.) are not write-protected. A user can overwrite them by assigning a new value to them.  For example, a statement such as “smua.measure.v = nil” will fail but “smua = nil” will not.  Once a Keithley-defined global variable is overwritten, the unit must be power cycled to restore normal operation.

Models Affected:

PR26902

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The display will briefly change to the source/measure screen during execution of a TSP-Link reset operation that has been initiated by front panel operation.

Models Affected:

PR26990

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

In prompting mode, a prompt is not always generated for an abort message if a tsplink.reset() command that was initiated from another command interface is executing when the abort messages is received.  The units will abort properly even though the prompt is not generated.

Models Affected:

PR26991

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

If a tsplink.reset() command is executing when a script (or command) is aborted, the tsplink.reset() command will be allowed to complete before execution is aborted.  The tsplink.reset() command may take several seconds when a large number of nodes are connected together.

Models Affected:

PR27182

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

There may be an expectation that the front panel store feature will utilize the reading buffer append mode, which can be set via a remote command interface.  It does not nor does it provide an option to append readings to the buffer. It only allows the user to overwrite the current buffer for each store operation.  The reading buffer’s append mode that can be set via a remote interface is ignored.

Models Affected:

PR27460

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Performing a tsplink.reset() for the first time after power-up of a system does not abort scripts running on nodes other than the one that initiated the reset.  This might allow multiple scripts to be running simultaneously.

Models Affected:

PR28889

PR32470

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

*CLS does not clear the output queue.

Models Affected:

PR29055

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The digital output port value displayed on the front panel may be incorrect after power-on.  If any of the digital I/O lines are externally driven low during power-on, the value displayed in the “GENERAL | DIGOUT | DIG-IO-OUTPUT” menu will reflect the state of the lines at power-up rather than their programmed output value.  After writing a value to the port via the front panel or a remote command, the correct value is displayed.

Models Affected:

PR29079

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Sending several thousand abort messages to the unit over the RS-232 interface while the unit is powering up can cause the unit to halt with an “UNKNOWN FATAL ERROR” message on the display.

Models Affected:

PR29249

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Using certain commands from a coroutine will cause the unit to become unresponsive, generate improper error messages, or fail to execute.  In addition to saving a script or listing a script the following commands will cause one or more of these problems when executed from a coroutine: delay(), makegetter(), makesetter(), print(), printbuffer(), printnumber(), script.delete(), script.new(), script.newautorun(), script.restore(), timer.clear(), and timer.wait().

Models Affected:

PR29951

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

While the unit is rebooting after performing a flash upgrade, any GPIB activity will cause the unit to halt with an “UNKNOWN FATAL EXCEPTION” error.  The firmware is updated correctly.  Cycling power on the unit corrects the problem.

Models Affected:

PR29994

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The pulse script will execute an smuX.reset() when aborted.  This will reset all the SMU settings to their defaults.

Models Affected:

PR32196

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Attempting to allocate a reading buffer when the unit is low on memory may cause an out of memory error when it should not.  When there is insufficient memory to allocate a reading buffer, the garbage collector should automatically run to reclaim any unused memory before giving up and generating an out of memory error.  This usually fails to happen.

Workaround:

To work around this issue, call collectgarbage() prior to creating a new reading buffer.

Models Affected:

PR32856

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The display.prompt() function may display some strange very small values instead of zero.  For example, you might see “-7.9409e-23” instead of “+0.0000e+00”.  This problem also affects factory script input routines because they use this function.

Models Affected:

PR32947

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Prior to performing a tsplink.reset(), the node table does not contain an entry for the local node under the local node index.  It only appears under the index “self.”  This also happens when the tsplink.reset() command fails with an error.

For example, consider a node with a node number of 5.  After a successful tsplink.reset(), the index 5 will refer to the local node.  Prior to executing the first tsplink.reset() or after a tsplink.reset() fails, the entry at index 5 in the node table is nil.

Models Affected:

PR32961

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

If you create a dynamic reading buffer on a remote node with the smuX.makebuffer() command and later perform a tsplink.reset() command, the reading buffer will be locked into memory with no way to access it or to delete it until the remote node is powered off.

Workaround:

This problem can be avoided by setting all dynamic reading buffer references to nil and calling collectgarbage() prior to executing tsplink.reset().

Models Affected:

PR33182

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Passing a line number other than 1 or 2 to the display.gettext() function will cause the instrument to become unresponsive.

Models Affected:

PR33183

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The error message for error 1103 reads “Max greater than min” when it should read “Min greater than max.”

Models Affected:

PR33303

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

ICL attributes that are integers can be assigned extremely large values without error.  The actual value assigned will be a truncated version of the large value the user is attempting to assign.

Models Affected:

PR33510

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

If hardware flow control is enabled on the RS-232 interface and the instrument is being controlled over that interface, unplugging the serial cable while a script is generating output will cause the instrument to hang when the output buffers fill up.  The EXIT button will not fully terminate the script until the cable is reconnected.  The instrument will recover if the cable is reconnected.

Models Affected:

PR33587

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

It is possible to abort the factory scripts by pressing the EXIT button or by sending an abort over the RS-232 interface during power-up when the instrument is running the factory scripts.  This prevents the factory scripts from creating some of the functions defined in them and prevents them from populating the front panel load menu.

Models Affected:

PR34254

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

The display annunciators can get out of synchronization with the state of the instrument and may fail to represent the true state of the instrument.  The instance of this problem most likely to be observed affects the REM indicator and can often be induced by sending a command message very quickly after sending an abort message.  This problem was introduced in version 1.2.0.

Models Affected:

PR34273

2635, 2636

Symptom:

Changing the state of the analog filter causes a measurement offset shift of approximately 50fA to 100fA.

Models Affected:

PR34493

PR34494

2601, 2602, 2611, 2612, 2635, 2636

Symptom:

Sending an abort message to the instrument when it is in local with the output on may cause the unit to stop updating the display with new measurements.  It may also generate a script interrupted error.

# v1.3.0 Release

## Overview

v1.3.0 was released to support the model 2635 and model 2636.  This is the initial firmware release for those models.  For the other models supported by this firmware, this release introduces three critical fixes, several enhancements, and several minor fixes.

## Compatibility Concerns

Version 1.3.0 firmware introduces changes that may affect compatibility with automated calibration software used to calibrate a Series 2600 instrument.  See enhancement PR31753 in the Enhancements section for more details.

## Critical Fixes

Models Affected:

PR31664

2611, 2612

Symptom:

After generating a current pulse greater than 1.5A using the KIPulse script, the SMU may output a large current after the pulse.  The magnitude of the current is proportional to (but larger than) the bias current specified for the pulse but even bias levels of zero may cause some current to be output.  When this happens, the SMU is capable of outputting currents in excess of 5A.  If the SMU is configured as a current source prior to initiating the pulse, the SMU will continue to output this current after the pulse script completes.  Large sustained currents caused by this issue may cause the unit to become unresponsive and possibly damage the SMU.

Resolution:

This issue has been corrected.

Models Affected:

PR32195

2601, 2602, 2611, 2612

Symptom:

An out of memory error while trying to allocate a reading buffer with the smuX.makebuffer command will crash the unit.  The unit must be power cycled to recover.

Resolution:

This issue has been corrected.

Models Affected:

PR32445

PR32956

2601, 2602, 2611, 2612

Symptom:

Autoranged measurements may return a bad measurement if the SMU needs to change ranges to make the measurement.  This can happen in the following scenarios:  It may return a bad measurement if it needs to downrange to the 100nA measure range while making the measurement.  It may return a bad measurement if it needs to downrage to or uprange from the lowest voltage measure range while making the measurement.  It may return a measurement made on a suboptimal range if it needs to uprange while making the measurement.

Resolution:

This issue has been corrected.

## Enhancements

Models Affected:

PR29810

2601, 2602, 2611, 2612

Enhancement:

The amount of time it takes to perform a TSP-Link® reset operation with 3 or more nodes has been substantially improved.  The execution time is now approximately one second per node.

Models Affected:

PR30789

PR30793

2611, 2612

Enhancement:

A new measurement autorange algorithm has been added that sets the measurement range to the range used for the source limit function.  This autorange algorithm is selected by setting the smuX.measure.autorangeY to the value smuX.AUTORANGE_FOLLOW_LIMIT.  Note: setting a fixed range will still set this attribute to smuX.AUTORANGE_OFF as usual regardless of which autorange algorithm is being used.

Models Affected:

PR30791

PR30844

2611, 2612

Enhancement:

Logic was added that is able to detect unstable output conditions that could damage the instrument.  When detected, the SMU will disable all low current measurement ranges to protect itself.  Although the SMU will continue to make measurements with an unstable output, the measurements may not be meaningful.

New status model bits were added for the unstable output condition.  The table below outlines the status model bits that have been added.  Note that the entire status.questionable.unstable_output register set has been added.

Register Set	Bit

status.questionable	UNSTABLE_OUTPUT (OU)

status.questionable.instrument.smua	UNSTABLE_OUTPUT (OU)

status.questionable.instrument.smub	UNSTABLE_OUTPUT (OU)

status.questionable.unstable_output	SMU A (SMUA)

status.questionable.unstable_output	SMU B (SMUB)

This enhancement was originally released in firmware version 1.1.4 for the 2601 and 2602 but it was not included in version 1.2.0 for the 2611 and 2612.  This release includes this enhancement for all products supported by this firmware.

Models Affected:

PR30211

PR30879

2601, 2602, 2611, 2612

Enhancement:

Two parameters were added to the ConfigPulseXMeasureY function in the KIPulse factory script immediately after the sync_out parameter.

The parameter sync_in_timeout was added to allow the timeout period for waiting for the sync_in trigger to be specified.  The timeout value is in seconds.  The sync_in_timeout parameter is optional and defaults to 10 seconds.

The sync_in_abort parameter was added to allow the user to specify if the pulse should be aborted when the input trigger is not received within the timeout period.  If this parameter is true the pulse will be aborted when no input trigger is detected within the timeout period.  The default value for this parameter is true.

Models Affected:

PR31753

2601, 2602, 2611, 2612

Enhancement:

A new attribute has been added to assist calibration labs in tracking calibration adjustments as well as calibration dates.  The smuX.cal.adjustdate attribute has been added.  Its usage is similar to the smuX.cal.date and smuX.cal.due attributes.

The smuX.cal.adjustdate attribute cannot be changed without first changing the calibration constants.  If the calibration constants are changed, they cannot be saved without first setting the smuX.cal.adjustdate attribute.  Setting the smuX.cal.date and smuX.cal.due are no longer required to be changed before saving the calibration data.  The smuX.cal.date and smuX.cal.due have no affect on the calibration constants so they may be changed and saved without making any changes to the calibration constants.

Models Affected:

PR32166

2601, 2602, 2611, 2612

Enhancement:

The KIPulse script functions have been enhanced to accept a table of values for the off_time parameter.  When a table of values is used for to specify the pulse off time(s), each pulse will use the appropriate off time from the table.  The first pulse will use the value at index 1, the second pulse will use the value from index 2, etc.  If a single value is given instead of a table or a table with a single value is given, that value will be used for all the pulses.  If there is more than one entry in the table, the number of entries must match the number of pulses.  When performing dual channel pulsing, the times for both channels must match.

Models Affected:

PR32167

2601, 2602, 2611, 2612

Enhancement:

The source level may now be changed during overlapped measurements as long as the source autorange setting is off.

Models Affected:

PR32577

2601, 2602, 2611, 2612

Enhancement:

The KIPulse script has been enhanced to allow simultaneous I and V measurements during pulsing.  To measure I and V simultaneously, pass a table of reading buffers to the ConfigPulseXMeasureY() function for the reading buffer parameter.  When building the table, store the reading buffer to hold the I measurements at index “i” and the reading buffer to hold the V measurements at index “v”.  For example:

status, msg = ConfigPulseVMeasureI(smua, 0, 10, 1e-3, 1e-3, 1e-3, 2, {i = smua.nvbuffer1, v = smua.nvbuffer2}, 1)

Models Affected:

PR33361

2601, 2602, 2611, 2612

Enhancement:

Two new attributes have been added that control measurement delays.  The smuX.measure.delay attribute can be set to a value in seconds that will cause the SMU to automatically wait that amount of time each time either smuX.measure.Y() or smuX.measure.overlappedY() is called.  If the measurement count is greater than one, the measurement delay will only be imposed before the first measurement.

This attribute may also be set to the value smuX.DELAY_AUTO.  When set to this value, the SMU will automatically choose a delay value based on the measurement range.  These automatic delay values can be systematically increased or decreased by setting the smuX.measure.delayfactor to a scaling factor.  The automatic delay values are multiplied by this scaling factor.  The default value for this attribute is 1.0.  Set this attribute to a value greater than one to increase the delays.  Set it to a value less than one to decrease the delays.  The smuX.measure.delayfactor attribute is ignored when the smuX.measure.delay attribute is not set to the value smuX.DELAY_AUTO.

Models Affected:

PR33362

2635, 2636

Enhancement:

The smuX.measure.analogfilter attribute can be set to the value 1 to enable an analog filter on the 1nA and 100pA ranges.  The analog filter will not be used on all other measurement ranges.  Set this attribute to zero to disable the analog filter.

Models Affected:

PR33543

2601, 2602, 2611, 2612, 2635, 2636

Enhancement:

A median filter can now be selected as a digital measurement filter.  From the remote interface, the median filter is selected by setting the smuX.measure.filter.type attribute to smuX.FILTER_MEDIAN.

When the median filter is enabled, measurements returned will be the median of the last smuX.measure.filter.count unfiltered measurement samples (sliding window).  If the sliding window is empty when a measurement is requested, the filter window will be pre-filled to capacity with copies of the first sample.

To accommodate selection of the median filter, the front panel filter configuration menus have changed slightly.  From the filter configuration menu, one may select TYPE or COUNT.  COUNT has the same meaning as the older AVERAGE-COUNT entry but the count now applies to both the average filter as well as the median filter.  Under TYPE, one may select AVERAGE or MEDIAN.  If AVERAGE is selected, the specific type of average filter is chosen by selecting either MOVING or REPEAT.

## Non-critical Fixes

Models Affected:

PR27088

2601, 2602, 2611, 2612

Symptom:

The tsplink.reset() command may cause one or more units on the TSP-Link network to become unresponsive.  This problem happens very infrequently.  The symptoms of this problem are identical to those of PR34098.  The likelihood of PR34098 occurring is very low compared to this problem.

Resolution:

This issue has been corrected.

Models Affected:

PR30054

2611, 2612

Symptom:

Front panel display updates cause timing jitter.  This is most apparent when making measurements that are stored in a reading buffer.

Workaround:

Turning off the display with the display.clear() command avoids these timing anomalies.

Resolution:

This issue has been corrected.

Models Affected:

PR30094

2611, 2612

Symptom:

When attempting to initiate a single channel pulse that does not make a measurement, the pulse script will abort with a run-time error.

Resolution:

This issue has been corrected.

Models Affected:

PR30129

2601, 2602, 2611, 2612

Symptom:

The buffer recall key will not cycle between all available reading buffers when the numeric keypad feature is enabled.  The buffer recall operation works properly when the numeric keypad feature is disabled.

Resolution:

This issue has been corrected.

Models Affected:

PR30247

PR30308

2601, 2602, 2611, 2612

Symptom:

The star annunciator on the front panel lights when storing readings to a reading buffer when initiated from the front panel but it does not when initiated via a remote command.

Resolution:

The star annunciator will light when an SMU is acquiring reading buffer data regardless of how the store operation was initiated.

Models Affected:

PR30248

PR30309

2601, 2602, 2611, 2612

Symptom:

Setting the display.smuX.digits attribute does not behave as documented in the manual.  The following table details the actual behavior when setting the attribute to the value X.  This behavior is incorrect.  Note that a value of 5.0 will select 4.5 digits, etc.

Value of X	Number of digits

X	<=	3.5	Error

3.5	<	X	<	5.5	4.5

5.5	<=	X	<	6.5	5.5

6.5	<=	X	<	7.5	6.5

7.5	<=	X			Error

Resolution:

This issue has been corrected.  The following table details the new and correct behavior.

Value of X	Number of digits

X	<=	3.0	Error

3.0	<	X	<	5.0	4.5

5.0	<=	X	<	6.0	5.5

6.0	<=	X	<	7.0	6.5

7.0	<=	X			Error

Models Affected:

PR30788

PR30792

2611, 2612

Symptom:

Passing parameters of the wrong datatype or with unexpected values to the ConfigXXX() function in the KIPulse script may cause the InitiatePulseTest() or InitiatePulseTestDual() function to generate errors when it is executed.  If this happens and the script is aborted, the SMU may not reset properly.

Resolution:

The ConfigXXX() functions now perform more thorough parameter checking and will return a failure indication and error message for problematic parameter values.

Models Affected:

PR31172

PR31182

2601, 2602, 2611, 2612

Symptom:

Attempting to set a fixed source range will turn the source-autorange setting off even if there is an error generated while processing the command.  When source-autorange is turned on, setting a new source range will automatically turn source-autorange off.  If there is an error generated for this operation, such as trying to set the new range lower than the source level, the command should not take any action.

Resolution:

This issue has been corrected.

Models Affected:

PR31483

2601, 2602, 2611, 2612

Symptom:

When a digital I/O line is configured for synchronous mode triggers, it will properly detect and latch a falling edge trigger.  Once latched, however, it cannot be released.

Resolution:

This issue has been corrected.

Models Affected:

PR31536

2611, 2612

Symptom:

Specifying a non-zero bias level with the KIPulse script does not work properly and causes the script to generate error messages when the pulse is generated.  If the bias is a positive value, the pulse will complete normally but the source range will not be restored.  If the bias is a negative value, the pulse will be generated properly but the bias level before the pulse will not be set.

Resolution:

This issue has been corrected.

Models Affected:

PR31749

2611, 2612

Symptom:

Sweeping from a negative value to zero using the KIPulse script result in a polarity error being generated.

Resolution:

This issue has been corrected.

Models Affected:

PR31749

2611, 2612

Symptom:

The KIPulse script allows logarithmic sweeps to start or end at zero.

Resolution:

This issue has been corrected.

Models Affected:

PR31775

2611, 2612

Symptom:

When using the KIPulse script, specifying a bias level higher than the pulse level will cause the InitiatePulseTest function to generate an error.

Resolution:

This issue has been corrected.

Models Affected:

PR31843

2611, 2612

Symptom:

The SMU may briefly output 0V when changing range to the 20V range (either manually or during autoranging) while the source is outputting a negative voltage.

Resolution:

This issue has been corrected.

Models Affected:

PR32510

2601, 2602, 2611, 2612

Symptom:

When the TSP-Link system is in the Local Control State (LCS), sending a command to the TSP-Link master node that will execute an SMU command on a remote TSP-Link node immediately after the LCS to Remote Control State (RCS) transition may cause the system to hang if any script or command is later aborted at any time.  This problem is more likely to be seen if the output on the remote SMU is on during the LCS to RCS transition.

Workaround:

This problem can be avoided by executing a 50 millisecond delay as the first command in a command message that causes an LCS to RCS transition.  For example, if the system is in the LCS state, send the message “delay(0.05)” as the first command to the master.  After this, all subsequent commands are safe until the system is again put in the LCS state.  As long as the system remains in the RCS state (no aborts are sent and the EXIT or ON/OFF buttons are not pressed), no further delays need be sent.

Resolution:

This issue has been corrected.

Models Affected:

PR32534

2601, 2602, 2611, 2612

Symptom:

Aborting a waitcomplete operation while a remote node is executing an overlapped operation will cause the TSP-Link system to hang.

Resolution:

This issue has been corrected.

Models Affected:

PR33493

PR34008

2601, 2602, 2611, 2612

Symptom:

If the display screen setting is changed from the front panel, the display.screen attribute accessible from the remote interface is not updated to reflect the new screen selection.

Resolution:

This issue has been corrected.

Models Affected:

PR33553

2611, 2612

Symptom:

Changing the source range from the 20V range to the 100V range while sourcing more than 20V will cause the output to glitch while changing range.

Resolution:

This issue has been corrected.

Models Affected:

PR33559

2611, 2612

Symptom:

The ConfigPulseIMeasureVSweepLin function from the KIPulse factory script will generate an error when configuring a negative sweep with a start value of 0.

Resolution:

This issue has been corrected.

Models Affected:

PR33670

2601, 2602, 2611, 2612

Symptom:

Changing source function from a 0A current source to a voltage source with the voltage source level lower than the voltage limit that was in effect as a current source may cause a large voltage spike on the output.

Resolution:

This issue has been corrected.

Models Affected:

PR33754

2601, 2602, 2611, 2612

Symptom:

If the output of a SMU is turned on and back off from the front panel and then the output is later turned on via a remote command interface, the display will display a stale measurement.  It will display the last measurement displayed when the output was on while in local rather than a stale measurement indication.

Resolution:

This issue has been corrected.

# v1.2.0 Release

## Overview

v1.2.0 was released to support the model 2611 and model 2612.  This is the initial firmware release for the model 2611 and model 2612.  This version was never approved for use on the model 2601 or model 2602.  All fixes and enhancements listed here are of interest for all models as these fixes and enhancements are included in all subsequent releases.

This release did not include the unstable output detection enhancement released under PR30149/PR30794 in version 1.1.4.

## Critical Fixes

There were no critical fixes included in this release.  See the Enhancements section and the Non-critical Fixes section for more information about release content.

## Enhancements

Models Affected:

PR27601

2601, 2602, 2611, 2612

Enhancement:

Added the Contact Check feature.  Documentation of the new commands and specifications for this feature is beyond the scope of this document and can be found in the product manual.

Models Affected:

PR28493
PR29980

2601, 2602, 2611, 2612

Enhancement:

Added the Line Frequency Auto-detect feature.  In addition to 50Hz or 60Hz options for the line frequency, AUTO can now be selected from the front panel.  From the remote interface, localnode.autolinefreq can be set to the value true to enable automatic line frequency detection.

When automatic line frequency detection is enabled, the unit will analyze environmental noise at power up and automatically select the best line frequency setting based on the noise measured.

Models Affected:

PR28643
PR29983

2601, 2602, 2611, 2612

Enhancement:

The default calibration password for all models has been changed to KI0026XX.  This avoids the confusion created by the model 2601 SMUs having different calibration passwords than the model 2602 SMUs.

Models Affected:

PR28862
PR28863

2601, 2602, 2611, 2612

Enhancement:

By default, the unit does not impose any measurement settling delays.  Changes in the output may cause transients in the measure circuit that are not settled out before making measurements.

A new attribute, smuX.source.delay, has been added to the command set.  The default value for this attribute is smuX.DELAY_OFF, which preserves the functionality of prior firmware versions.  If this attribute is set to smuX.DELAY_AUTO, the SMU will impose a settling delay based on the i-measure range.  Setting this attribute to any positive value will cause the SMU to always impose a delay equal to that value regardless of range.

Models Affected:

PR29228

2601, 2602, 2611, 2612

Enhancement:

A new command has been introduced to read the state of the 2611/2612 interlock signal.  The smuX.readinterlock() function will return true if the interlock signal is asserted, and false if it is not.

This function is present on the 2601/2602 but will return the state of the output enable signal.

Models Affected:

PR29479
PR29901

2601, 2602, 2611, 2612

Enhancement:

Numeric keypad entry can now be disabled.  When the display is in numeric entry mode, some buttons such as Channel A SRC, DIGITS, etc. have a different meaning.  Disabling the numeric keypad entry feature prevents this from causing confusion.

The numeric keypad feature can be controlled with the display.numpad attribute.  Set this attribute to display.ENABLE to enable the numeric keypad feature.  Set it to display.DISABLE to disable it.  For example, send the command “display.numpad = display.DISABLE” to disable the numeric keypad feature.

The numeric keypad feature can also be enabled or disabled from the front panel via the MENU | GENERAL | DISPLAY menu.

Models Affected:

PR29684
PR29695

2601, 2602, 2611, 2612

Enhancement:

All ASCII characters below 0x20 (space) are treated as space characters except for the following, which are handled the same as they were in earlier versions of firmware:

0x08 (BS)	The backspace character will remove the previous character of a message.  This is useful for typing messages directly from an RS-232 terminal or terminal application.

0x09 (TAB)	The tab character will retain its value because it has meaning for formatting output.

0x0A (LF)	The linefeed character is used as a message terminator as it is in IEEE-488.2.  It is also used to separate lines in a multi-line script.

This change provides more compatibility with the IEEE-488.2 standard and avoids syntax errors caused by ASCII control characters being embedded in messages sent to the unit.

Models Affected:

PR29803

2601, 2602, 2611, 2612

Enhancement:

The unit will display the saved setup number being loaded at power-on if one (other than the default factory setup) is being loaded.

## Non-critical Fixes

Models Affected:

PR27835
PR29991

2601, 2602

Symptom:

Fatal errors may not display on the front panel while the unit is running a CPU intensive script.  The errors will eventually display when the script performs a delay or some other action that yields the CPU.

Resolution:

This issue has been corrected.

Models Affected:

PR27939
PR29367

2601, 2602

Symptom:

The unit will allow values other than 1 or 0 to be used to set the collectsourcevalues and collecttimestamps attributes of a reading buffer.

Resolution:

The unit will now generate an illegal parameter error when trying to set these reading buffer attributes to a value other than 0 or 1.

Models Affected:

PR28176
PR29982

2601, 2602

Symptom:

When turning the output off while sourcing voltage with a compliance limit above 100mA, the unit may occasionally use a fast shutdown sequence that causes glitches on the output.

Resolution:

This issue has been corrected.

Models Affected:

PR28720
PR29984

2601, 2602

Symptom:

Fatal errors generated when the unit is booting are not always displayed.  When this happens, they will appear after pressing a key.

Resolution:

This issue has been corrected.

Models Affected:

PR29058

2601

Symptom:

Resetting the calibration password on a model 2601 by shorting the cal pads causes an error regarding “smub” to be generated.  The error can be safely ignored.  This issue does not affect the model 2602.

Resolution:

This issue has been corrected.

Models Affected:

PR29286
PR29291
PR29366

2601, 2602

Symptom:

The attributes of a nonvolatile reading buffer are set to their defaults after a power cycle when the reading buffer is empty.

Resolution:

Nonvolatile reading buffers now retain their attribute settings across power cycles even when a buffer is empty.

Models Affected:

PR29289
PR29290

2601, 2602

Symptom:

The beeper enable/disable setting is not saved or restored with saved setups.

Resolution:

This issue has been corrected.

Models Affected:

PR29318
PR29319

2601, 2602

Symptom:

When changing the source lowrange setting of an SMU to a value above the range being used while the source is configured for autoranging, the range may not be changed until the source level of the SMU is changed.

Resolution:

This issue has been corrected.

Models Affected:

PR29320
PR29321

2601, 2602

Symptom:

When using the CONFIG button on the front panel, the unit intermittently ignores all configuration sequence button presses and stays on the configuration menu.

Resolution:

This issue has been corrected.

Models Affected:

PR29774
PR29778
PR29985

2601, 2602

Symptom:

The RQS bit in the serial poll byte reported during a GPIB spoll operation is not set when the unit is asserting the SRQ line.  The unit will properly assert the SRQ line but will not report this in the RQS bit during a GPIB spoll operation.  This problem was introduced in v1.1.1.

Resolution:

This issue has been corrected.

Models Affected:

PR29786

2601, 2602

Symptom:

When sourcing current, the unit will display the wrong source range and level after the current source low range setting is changed.  The values are displayed correctly again after setting a new output level or selecting a new source range.

Resolution:

This issue has been corrected.

Models Affected:

PR29827
PR29867

2601, 2602

Symptom:

When trying to access an element at index zero in a reading buffer, the error message will refer to an index of “(null)”.

Resolution:

This issue has been corrected.

# v1.1.4 Release

## Overview

v1.1.4 was a maintenance version of the 2601/2602 firmware but was never released.

## Critical Fixes

Symptom:

PR30439

PR30855

PR30856

When the unit encounters a GPIB error, a 5046 “SMU too hot” error, or an 802 “OUTPUT blocked by interlock” error the unit may become unstable or unresponsive.  This problem was introduced in v1.1.0 of the firmware.

Resolution:

This issue has been corrected.

## Enhancements

Enhancement:

PR30149

PR30794

Logic was added that is able to detect unstable output conditions that could damage the instrument.  When detected, the SMU will disable all low current measurement ranges to protect itself.  Although the SMU will continue to make measurements with an unstable output, the measurements may not be meaningful.

New status model bits were added for the unstable output condition.  The table below outlines the status model bits that have been added.  Note that the entire status.questionable.unstable_output register set has been added.

Register Set	Bit

status.questionable	UNSTABLE_OUTPUT (OU)

status.questionable.instrument.smua	UNSTABLE_OUTPUT (OU)

status.questionable.instrument.smub	UNSTABLE_OUTPUT (OU)

status.questionable.unstable_output	SMU A (SMUA)

status.questionable.unstable_output	SMU B (SMUB)

## Non-critical Fixes

Symptom:

PR30455

Some error messages are missing words when they are displayed on the front panel.  Only error messages that are split between multiple display pages are affected.  The message text in the error queue is unaffected by this problem.

Resolution:

This issue has been corrected.

# v1.1.3 Release

## Overview

v1.1.3 was a maintenance release of the 2601/2602 firmware.

## Critical Fixes

Symptom:

PR29137
PR29262
PR29654

Setting the voltage or current limit to a negative value in most circumstances will cause the unit to become unresponsive requiring a power cycle.

Resolution:

This issue has been corrected.

Symptom:

PR29418
PR29655
PR29656

Setting the appendmode, collectsourcevalues, and collecttimestamps attributes of a reading buffer for an SMU on a remote node does not work.  When attempting to set one of these attributes on a remote node, either an error will be generated or the attribute will be set to a semi-random value.  This problem does not affect reading buffers on the local node.

Resolution:

This issue has been corrected.

## Non-critical Fixes

Symptom:

PR29088
PR29657
PR29658

The RS-232 interface does not operate at baud rates below 9600 baud.

Resolution:

This issue has been corrected.

# v1.1.2 Release

## Overview

v1.1.2 was a maintenance release of the 2601/2602 firmware issued strictly to support manufacturability of the product.

# v1.1.1 Release

## Overview

v1.1.1 was a maintenance release of the 2601/2602 firmware.  It addresses a few critical issues, offers a few minor product enhancements, and corrects several non-critical issues

## Critical Fixes

Symptom:

PR27680

The first measurement after a measurement range change from the 1A range to a lower range may be incorrect.  This is due to SMU circuit settling and may affect any measurement taken within 1mS of the range change.  This affects manual range changes as well as range changes initiated by an auto-range measurement.

Resolution:

The SMU imposes longer internal settling delays when down-ranging the measure circuit from the 1A range.

Symptom:

PR28026

When changing the source value to the opposite polarity, the compliance circuit is not updated properly causing the SMU compliance level to be slightly off from the user specified value.  Setting a new limit value updates the circuit settings and restores proper behavior.

Resolution:

This issue has been corrected.

Symptom:

PR28097

If the SMU is sourcing voltage, a load change or source value change causes more than 100mA of current to flow, the current limit is set to a value greater than 100mA, the measure range is less than 1A, and a measurement is made within 500uS of the current going above 100mA, then the SMU may return a bad current measurement.

Resolution:

This issue has been corrected.

## Enhancements

Enhancement:

PR26966

SMU errors now include an indication of which SMU channel was involved in generating the error.

Enhancement:

PR27752

A new function has been added to the command set that returns the available amount of dynamically allocated memory.  The meminfo() function returns the amount of available dynamically allocated memory and the total amount of dynamic memory in the product.  The new command is used as follows:

freemem, totalmem = meminfo()

The first return value, freemem, is the amount of memory still available for dynamic memory allocation and is given in kilobytes.  The second return value, totalmem, is the total amount of dynamic memory built into the product and is also given in kilobytes.

Dynamically allocated memory is used to store the run-time environment, user created reading buffers, loaded scripts, and for general firmware operation.

## Non-critical Fixes

Symptom:

PR27160

When setting the NPLC aperture, the unit will adjust the setting to the closest aperture the instrument can support.  When the value is read back, it should reflect the value chosen by the instrument; however the instrument reports the exact value specified by the user and not the value actually being used.

Resolution:

When reading back the NPLC aperture, the actual aperture being used by the SMU will be reported.

Symptom:

PR27608

Run-time errors detected while executing a script with loadandrunscript are not reported.  Syntax errors are reported correctly but run-time errors detected while running the script as part of loading it are not.  All run-time errors are reported correctly if the script is executed again after it is loaded.

Resolution:

This issue has been corrected.

Symptom:

PR27628

The unit will allow values other than 1 or 0 to be used to set the append mode of a reading buffer.

Resolution:

The unit will now generate an illegal parameter error when trying to set a reading buffer append mode attribute to a value other than 0 or 1.

Symptom:

PR27645

The status byte reported for a GPIB serial poll is only updated when the unit’s SRQ state changes.  This causes old status byte values to be reported for GPIB spoll operations.  This does not affect reading the status byte via other remote commands.

Resolution:

This issue has been corrected.

Symptom:

PR27653

Some TSP-Link initialization failure messages are reported up to ten times on remote units when the TSP-Link is reset and initialization failures are detected.

Resolution:

This issue has been corrected.

Symptom:

PR27718

If smuX.makebuffer() is called with a buffer size of zero, the unit will generate an out of memory error rather than the expected illegal parameter value error.

Resolution:

The unit will now generate an illegal parameter value error in this situation.

Symptom:

PR27771

If back to back command messages are sent over the GPIB interface fast enough, some characters at the beginning of the second command message may be lost.  There are no GPIB controllers known to be able to run this fast.  However, this can happen when fooling the GPIB controller into sending two messages faster by performing one send operation that contains linefeed characters in the middle of the data.  The 26XX will properly interpret the linefeeds (without EOI) as message termination characters thereby interpreting the one send as multiple command messages.

Resolution:

The GPIB interface now holds off bus communication on receipt of a newline character long enough for the unit to recognize the receipt of a complete command message.  This fix does not degrade communication throughput.

Symptom:

PR27794

TSP-Link initialization failures can cause all units on the TSP-Link to become unresponsive.  Setting two nodes to the same TSP-Link address will exhibit this problem.

Resolution:

This issue has been corrected.

Symptom:

PR27859

Some pairs of units have an unusually high packet corruption rate on the TSP-Link.  The TSP-Link protocol automatically recovers from these problems but the issue can be noticed as degradation in performance when controlling remote units over the TSP-Link.

Resolution:

This issue has been corrected.  This fix corrects a synchronization problem eliminating the root cause of the packet corruption.

Symptom:

PR27968

PR28021

When using the SMU as a current source after using the High-Z output off mode, the SMU will report that it is in compliance when it is not.

Resolution:

This issue has been corrected.

# v1.1.0 Release

## Overview

v1.1.0 was an internal release only.  It was never released to the field.

# v1.0.X Releases

## Overview

The history of these releases is not included in this release notes document.  Please see the release notes from previous releases for more information about these earlier releases.

Recalibration is not necessary when upgrading to v1.3.3.  However, if the unit is recalibrated after performing the upgrade, a recalibration will be required if the firmware is later downgraded to any earlier version.

Due to the comparatively faster test script execution of this release, production test suites should be re-qualified to assure there are no DUT settling issues.

This version of firmware introduces some changes in behavior as compared to older firmware.  See the Compatibility Concerns section for more details.

Review the list of changes made in all versions in between your current version and this version.  Upgrade if any of the fixes or enhancements are desired.

These versions of firmware should be upgraded to v1.0.5, v1.1.3, or v1.3.3.

Version 1.1.5 and 1.1.6 were special versions of firmware that were not released to the general public.  Version 1.1.5 firmware should be upgraded to v1.1.6, or v1.3.3.