## Sheet1
Calculators based on Maynuo published specifications. Enter any value in the yellow boxes to derive Min and Max
| M9711 / M9712 Specs |  |  |  | Min | Max |  |  | M9710 Specs |  |  |  | Min | Max
| 3 | amp range |  | 3 |  |  |  |  | 3 | amp range |  | 3
| 0.03% | measurement |  | 0.0009 |  |  |  |  | 0.05% | measurement |  | 0.0015
| 0.05% | full scale |  | 0.0015 |  |  |  |  | 0.01% | full scale |  | 0.0003
|  |  |  | 0.0024 | 2.9976 | 3.0024 | [5 empty] | 0.0018 | 2.9982 | 3.0018
| 30 | amp range |  | 10 |  |  |  |  | 30 | amp range |  | 10
| 0.03% | measurement |  | 0.003 |  |  |  |  | 0.05% | measurement |  | 0.005
| 0.08% | full scale |  | 0.024 |  |  |  |  | 0.01% | full scale |  | 0.003
|  |  |  | 0.027 | 9.973 | 10.027 | [5 empty] | 0.008 | 9.992 | 10.008
| 20 | volt range |  | 3 |  |  |  |  | 20 | volt range |  | 3
| 0.015% | measurement |  | 0.00045 |  |  |  |  | 0.015% | measurement |  | 0.00045
| 0.03% | full scale |  | 0.006 |  |  |  |  | 0.05% | full scale |  | 0.01
|  |  |  | 0.00645 | 2.9935 | 3.0065 | [5 empty] | 0.01045 | 2.9895 | 3.0105
| 150 | volt range |  | 150 |  |  |  |  | 150 | volt range |  | 150
| 0.015% | measurement |  | 0.0225 |  |  |  |  | 0.015% | measurement |  | 0.0225
| 0.03% | full scale |  | 0.045 |  |  |  |  | 0.05% | full scale |  | 0.075
|  |  |  | 0.0675 | 149.9325 | 150.0675 | [5 empty] | 0.0975 | 149.9025 | 150.0975
| Front Panel measurements |  |  | 20 | °C | 51% | RH | 2016-04-03 |  | M9710
| PSU: |  | Rigol DP832
| Voltmeter: |  | Keithley 2000
| Ammeter: |  | Rigol DM3058E
All voltage readings taken without any load. Ampere readings taken by dialing in the displayed Maynuo load while observing the DMM to get as close as possible to test example.
[8 empty] | M9710 calculated specs |  |  | M9711/12 calculated specs |  |  | Official provided Cal. Cert. |  | Note
| 3 Ampere Range |  |  | DMM |  | Displayed | Error | Min | Max |  | Min | Max |  | Min | Max
| 0.1A |  |  | 0.10002 |  | 0.0997 | 0.321% | 0.0997 | 0.1004 |  | 0.0985 | 0.1015 |  | 0.0985 | 0.1015 | 1
| 1.5A |  |  | 1.5 |  | 1.4995 | 0.03134% | 1.499 | 1.5011 |  | 1.4981 | 1.502 |  | 1.498 | 1.502
| 3A |  |  | 3.0001 |  | 2.9993 | 0.02667% | 2.9982 | 3.0018 |  | 2.9976 | 3.0024 |  | 2.9975 | 3.0025
30 Ampere Range
| 1A |  |  | 1.0002 |  | 1 | 0.02% | 0.997 | 1.004 |  | 0.976 | 1.024 |  | 0.985 | 1.015 | 2
| 15A |  |  | n/a |  | n/a |  | 14.99 | 15.011 |  | 14.972 | 15.029 |  | 14.98 | 15.02
| 30A |  |  | n/a |  | n/a |  | 29.982 | 30.018 |  | 29.967 | 30.033 |  | 29.975 | 30.025
20V Range, Local
| 3V |  |  | 3.0004 |  | 3.009 | -0.2871% | 2.99 | 3.01 |  | 2.994 | 3.006 |  | 2.996 | 3.004 | 3
| 10V |  |  | 10.0002 |  | 10.008 | -0.07794% | 9.989 | 10.012 |  | 9.993 | 10.008 |  | 9.994 | 10.006
| 20V |  |  | 20.0002 |  | 20.005 | -0.02399% | 19.987 | 20.013 |  | 19.991 | 20.009 |  | 19.992 | 20.008
150V Range, Local
| 3V |  |  | 3.0002 |  | 3.01 | -0.3262% | 2.92 | 3.08 |  | 2.95 | 3.05 |  | 2.97 | 3.03
| 60V |  |  | 60.0005 |  | 59.99 | 0.0175% | 59.92 | 60.08 |  | 59.95 | 60.05 |  | 59.95 | 60.05
| 150V |  |  | n/a |  | n/a |  | 149.9 | 150.1 |  | 149.93 | 150.07 |  | 149.93 | 150.07
20V Range, Remote
| 3V |  |  | 3.0003 |  | 2.997 | 0.1114% | 2.99 | 3.01 |  | 2.994 | 3.006 |  | 2.996 | 3.004
| 10V |  |  | 10.0002 |  | 9.988 | 0.1221% | 9.989 | 10.012 |  | 9.993 | 10.008 |  | 9.994 | 10.006 | 4
| 20V |  |  | 20.0002 |  | 19.975 | 0.1262% | 19.987 | 20.013 |  | 19.991 | 20.009 |  | 19.992 | 20.008
150V Range, Remote
| 3V |  |  | 3.0002 |  | 3 | 0.005333% | 2.92 | 3.08 |  | 2.95 | 3.05 |  | 2.97 | 3.03
| 60V |  |  | 60.0006 |  | 59.92 | 0.1345% | 59.92 | 60.08 |  | 59.95 | 60.05 |  | 59.95 | 60.05 | 5
| 150V |  |  | n/a |  | n/a |  | 149.9 | 150.1 |  | 149.93 | 150.07 |  | 149.93 | 150.07
Notes
1 | Factory calibration "certificate" min/max is within a least significant digit of the derived values from published specs for M9711/12.
2 | For some reason the "certificate" appears to be using 0.05% of full scale instead of the published 0.08% in the high amp mode. Of course, the M9710 is consistent with its superior 0.01% if full scale compared to other models.
Overall, the M9710 is well within specifications on current readings, and clearly within its own much superior specs to the M9711/12. ;)
3 | Here the front panel low volt range, the M9710 fails the M9711/12 specs. Not sure why the specs on the cert are so much tighter than the calculated specs. However, it is still within the published M9710 specs.
4 | Now the M9710 doesn't meet it's own specs with remote voltage. I could overlook the 10V but the 20V is pushing it…
5 | Right on the edge of the calculated specs, but a fail based on M9711/12 specs and the M9710 "certificate" that appears to use the 9711/12 specs anyway.