## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Hewlett-Packard |  |  | S/N |  | Ref2
Model Number |  | HP3458A |  |  | Ambient Temp. |  | 26.5
Meter Info |  | HP3458A STD |  |  | Test date |  | 19-03-30 16:14:51
Options |  | 0,0 |  |  | Test Type |  | verify
Firmware |  | 9,2 |  |  | Notes |  | Test Front Ports
Last ext. calibration date |  | 03/23/2019 |  |  | 3458A NPLC? |  | 100
DUT CAL number? |  | 183 |  |  | DUT Internal TEMP? |  | 36.1
CAL? 58 Temp of CAL 0 |  | 34.7126710E+00 |  |  | CAL?1,1 |  | 39.9992704E+03
CAL?59 Temp of CAL 10 |  | 33.6766328E+00 |  |  | CAL?72 |  | 982.329424E-03
CAL? 60 Temp of CAL 10K |  | 36.4117225E+00 |  |  | CAL?2,1 |  | 7.07033766E+00
DESTRUCTIVE OVERLOADS Before |  | 304, DESTRUCTIVE OVERLOADS valid 2941 |  |  | DESTRUCTIVE OVERLOADS After |  | 304, DESTRUCTIVE OVERLOADS valid 2941
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
|  |  |  | -- | -- |  | --
Fluke 792X |  | 9.9999800 VDC |  | 2.2 ppm | 2018-02-01 |  | 2019-02-01
Fluke SL935 |  | 1.00006002 Ohm |  | 0.17 ppm | 2018-05-30 |  | 2019-05-30
Fluke SL935 |  | 9999.9760 Ohm |  | 0.33 ppm | 2018-05-30 |  | 2019-05-30
MFC
MFC Calibrate Date |  | 021019 |  |  | MFC last calibrated |  | 48 days
Calibrate Date WBFLAT |  | 020919 |  |  | MFC since WBFLAT |  | 49 days
Calibrate Date WB Gain |  | 021019 |  |  | MFC since WBGAIN |  | 48 days
MFC Calibrate Date Zero |  | 033019 |  |  | MFC since DCV Zero |  | 0 days
MFC Confidence level |  | 5720A 99% 24 Hours |  |  | CAL CONST 11V Range +zero |  | 398.1803
CAL CONST 6.5V |  | +6.957480629534423E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1793
CAL CONST 13V reference voltage |  | +1.385529947933807E+01 |  |  | CAL CONST, Zero calibration temperature |  | 24.6
CAL CONST 10KOHM true output resistance |  | +9.999785888249386E+03 |  |  | CAL CONST, ALL calibration temp |  | 24.6
CAL SONST 10kOHM standard resistance |  | +9.998741377818425E+03 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 25
CAL CONST 1 OHM |  | +9.998156865545850E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 227 days 18.20 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 3 days 14.08 hours |  |  | CAL CONST,DACLIN |  | -0.207
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT 24h Spec uV | Test Result % of SPEC | MEAS SDEV μV
0 .0mVDC | 0 | 0.43392 | 0.75 | -0.91 | 0.91 | N/A | 0.16 | 47.6839 | 0.010719 | +4.0000E-07,V,1
0.0VDC | 0 | 0.43605 | 0.75 | -0.9 | 0.9 | N/A | 0.15 | 48.4495 | 0.023835 | +4.0000E-07,V,1
00.0VDC | 0 | 1.0331 | 0.75 | -1.07 | 1.07 | N/A | 0.32 | 96.5542 | 0.070103 | +4.0000E-07,V,1
000.0VDC | 0 | 29.4229 | 0.75 | -14.75 | 14.75 | N/A | 14 | 199.4775 | 3.844 | +4.0000E-07,V,1
0000.0VDC | 0 | 84.0599 | 0.75 | -41.75 | 41.75 | N/A | 41 | 201.341 | 10.7757 | +4.0000E-07,V,1
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT 24h Spec ppm | Test Result % of SPEC | MEAS SDEV μV
0.01 | 100mV | 0.01 | 55 | 0.0099991 | 0.010001 | 10.636 | 32.5 | 12.1554 | 1.9117 | +4.4000E-07,V,1
0.02 | 100mV | 0.02 | 30 | 0.019999 | 0.020001 | 7.3185 | 17.5 | 15.4074 | 1.3493 | +4.8000E-07,V,1
0.05 | 100mV | 0.05 | 15 | 0.049999 | 0.050001 | 4.7376 | 8.5 | 20.16 | 0.34361 | +1.2000E+01,PPM,1
0.1 | 100mV | 0.1 | 10 | 0.099998 | 0.1 | 3.59 | 5.5 | 23.1613 | 0.10417 | +8.0000E+00,PPM,1
-0.01 | 100mV | -0.0099999 | 55 | -0.0099991 | -0.010001 | -5.219 | 32.5 | -5.9646 | -2.5914 | +4.4000E-07,V,1
-0.02 | 100mV | -0.02 | 30 | -0.019999 | -0.020001 | -2.748 | 17.5 | -5.7853 | -1.2553 | +4.8000E-07,V,1
-0.05 | 100mV | -0.05 | 15 | -0.049999 | -0.050001 | -0.0328 | 8.5 | -0.13957 | -0.31858 | +1.2000E+01,PPM,1
-0.1 | 100mV | -0.1 | 10 | -0.099998 | -0.1 | 1.148 | 5.5 | 7.4065 | -0.20157 | +8.0000E+00,PPM,1
0.1 | 1V | 0.1 | 10 | 0.099999 | 0.1 | 4.879 | 4.5 | 33.6483 | 0.33069 | +8.0000E+00,PPM,1
0.2 | 1V | 0.2 | 7.5 | 0.2 | 0.2 | 3.968 | 3 | 37.7905 | 0.12504 | +6.0000E+00,PPM,1
0.5 | 1V | 0.5 | 5.1 | 0.5 | 0.5 | 1.758 | 2.1 | 24.4167 | 0.26094 | +4.4000E+00,PPM,1
1 | 1V | 1 | 4.3 | 0.99999 | 1 | 1.93 | 1.8 | 31.6393 | 0.053983 | +3.7000E+00,PPM,1
-0.1 | 1V | -0.1 | 11.5 | -0.099998 | -0.1 | 1.044 | 4.5 | 6.525 | -0.25826 | +8.0000E+00,PPM,1
-0.2 | 1V | -0.2 | 7.5 | -0.2 | -0.2 | 1.928 | 3 | 18.3619 | -0.15751 | +6.0000E+00,PPM,1
-0.5 | 1V | -0.5 | 5.1 | -0.5 | -0.5 | 1.9646 | 2.1 | 27.2861 | -0.14169 | +4.4000E+00,PPM,1
-1 | 1V | -1 | 4.3 | -0.99999 | -1 | 1.756 | 1.8 | 28.7869 | -0.02637 | +3.7000E+00,PPM,1
1 | 10V | 1 | 4.3 | 0.99999 | 1 | 2.373 | 1 | 44.7736 | 0.097488 | +3.7000E+00,PPM,1
2 | 10V | 2 | 3.9 | 2 | 2 | 1.637 | 0.75 | 35.2043 | 0.043277 | +3.3500E+00,PPM,1
5 | 10V | 5 | 3.1 | 5 | 5 | 1.8158 | 0.6 | 49.0757 | 0.054768 | +2.5000E+00,PPM,1
10 | 10V | 10 | 2.8 | 10 | 10 | 1.464 | 0.55 | 43.7015 | 0.034007 | +2.2500E+00,PPM,1
-1 | 10V | -1 | 4.3 | -0.99999 | -1 | 0.832 | 1 | 15.6981 | -0.12316 | +3.7000E+00,PPM,1
-2 | 10V | -2 | 3.9 | -2 | -2 | 0.8755 | 0.75 | 18.828 | -0.030699 | +3.3500E+00,PPM,1
-5 | 10V | -5 | 3.1 | -5 | -5 | 1.2414 | 0.6 | 33.5514 | -0.06263 | +2.5000E+00,PPM,1
-10 | 10V | -10 | 2.8 | -10 | -10 | 1.07 | 0.55 | 31.9403 | -0.026079 | +2.2500E+00,PPM,1
10 | 100V | 10 | 2.8 | 9.9999 | 10.0001 | 4.2 | 5.5 | 50.6024 | 0.58882 | +2.2500E+00,PPM,1
20 | 100V | 20.0001 | 2.65 | 19.9999 | 20.0001 | 2.65 | 4 | 39.8496 | 0.235 | +2.2000E+00,PPM,1
50 | 100V | 50.0001 | 4.5 | 49.9996 | 50.0004 | 2.5852 | 3.1 | 34.0158 | 0.12552 | +3.8000E+00,PPM,1
100 | 100V | 100.0002 | 4 | 99.9993 | 100.0007 | 1.93 | 2.8 | 28.3824 | 0.052998 | +3.4000E+00,PPM,1
-10 | 100V | -10 | 2.8 | -9.9999 | -10.0001 | 0.627 | 5.5 | 7.5542 | -0.45603 | +2.2500E+00,PPM,1
-20 | 100V | -20 | 2.65 | -19.9999 | -20.0001 | 1.24 | 4 | 18.6466 | -0.097882 | +2.2000E+00,PPM,1
-50 | 100V | -50.0002 | 4.5 | -49.9996 | -50.0004 | 3.2612 | 3.1 | 42.9105 | -0.2074 | +3.8000E+00,PPM,1
-100 | 100V | -100.0002 | 4 | -99.9993 | -100.0007 | 2.181 | 2.8 | 32.0735 | -0.036927 | +3.4000E+00,PPM,1
-100 | 1000V | -100.0001 | 4 | -99.9993 | -100.0007 | 1.122 | 3.5 | 14.96 | -0.06115 | +3.4000E+00,PPM,1
-200 | 1000V | -200.0001 | 3.75 | -199.9986 | -200.0014 | 0.7285 | 3.48 | 10.0761 | -0.06381 | +3.2000E+00,PPM,1
-500 | 1000V | -500.0023 | 6 | -499.9941 | -500.0059 | 4.5094 | 5.7 | 38.5419 | -0.042591 | +4.8000E+00,PPM,1
-1000 | 1000V | -1000.0048 | 5.5 | -999.9799 | -1000.0201 | 4.816 | 14.6 | 23.9602 | -0.3149 | +4.4000E+00,PPM,1
100 | 1000V | 100.0002 | 10 | 99.9986 | 100.0014 | 2.156 | 3.5 | 15.9704 | 0.088239 | +3.4000E+00,PPM,1
200 | 1000V | 200.0003 | 7.5 | 199.9978 | 200.0022 | 1.5865 | 3.48 | 14.449 | 0.044328 | +3.2000E+00,PPM,1
500 | 1000V | 500.0025 | 6 | 499.9941 | 500.0059 | 5.0278 | 5.7 | 42.9726 | 0.13303 | +4.8000E+00,PPM,1
1000 | 1000V | 1000.0049 | 5.5 | 999.9799 | 1000.0201 | 4.914 | 14.6 | 24.4478 | 0.24586 | +4.4000E+00,PPM,1
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | 24h Spec ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99982 | 0.99978 | 85 | 0.9997 | 0.99994 | -31.4629 | 34.9991 | -26.2193 | 10.5844 | +7.0000E+01,PPM,1
1.9 Ω | 1.8999 | 1.8998 | 85 | 1.8997 | 1.9001 | -44.6421 | 20.7892 | -101.9479 | 2.538 | +7.0000E+01,PPM,1
10 Ω | 10.0003 | 10.0003 | 23 | 10 | 10.0006 | -6.0898 | 8.0002 | -19.6444 | 0.51897 | +2.0000E+01,PPM,1
19 Ω | 19 | 19 | 23 | 18.9992 | 19.0008 | -2.6753 | 18.7895 | -9.2925 | 2.825 | +2.0000E+01,PPM,1
100 Ω | 100.0032 | 100.0032 | 10 | 100.0016 | 100.0048 | 0.57398 | 6.0001 | 3.5874 | 0.46931 | +8.0000E+00,PPM,1
190 Ω | 189.9982 | 189.9982 | 10 | 189.9939 | 190.0025 | 0.0031579 | 12.5263 | 0.015385 | 0.23151 | +8.0000E+00,PPM,1
1 kΩ | 1000.0099 | 1000.0095 | 8 | 999.9997 | 1000.0201 | -0.447 | 2.2 | -4.3823 | 0.058107 | +6.5000E+00,PPM,1
1.9 kΩ | 1900.0239 | 1900.0222 | 8 | 1900.0029 | 1900.0449 | -0.9042 | 3.0527 | -8.1808 | 0.2771 | +6.5000E+00,PPM,1
10 kΩ | 9999.786 | 9999.7542 | 8 | 9999.684 | 9999.888 | -3.1801 | 2.2 | -31.1773 | 0.059734 | +6.5000E+00,PPM,1
19 kΩ | 18999.392 | 18999.371 | 9 | 18999.163 | 18999.621 | -1.1053 | 3.0526 | -9.1706 | 0.13868 | +7.5000E+00,PPM,1
100 kΩ | 99994.64 | 99993.5882 | 9 | 99993.5201 | 99995.7599 | -10.5183 | 2.1999 | -93.914 | 0.10838 | +7.5000E+00,PPM,1
190 kΩ | 189988.81 | 189985.9082 | 9 | 189984.2004 | 189993.4196 | -15.2735 | 15.2626 | -48.8556 | 0.12995 | +7.5000E+00,PPM,1
1 MΩ | 999980.6 | 999891.5895 | 16 | 999953.6007 | 1000007.5993 | -89.0122 | 10.9998 | -329.6773 | 0.086781 | +1.3000E+01,PPM,1
1.9 MΩ | 1899967 | 1899685.929 | 17 | 1899789.7047 | 1900144.2953 | -147.9347 | 76.3149 | -135.3289 | 0.5199 | +1.4000E+01,PPM,1
10MΩ | 9999029 | 9990677.408 | 33 | 9998149.134 | 9999908.866 | -835.2403 | 54.9951 | -949.1891 | 1.0082 | +2.7000E+01,PPM,1
19 MΩ | 18998410 | 18971204.2 | 43 | 18987094.742 | 19009725.258 | -1432.004 | 552.5897 | -219.434 | 2.5713 | +3.5000E+01,PPM,1
100 MΩ(2W) | 100007070 | 99988404.96 | 100 | 99946062.1521 | 100068077.8479 | -186.6372 | 510.0353 | -30.5945 | 1.4998 | +8.5000E+01,PPM,1
19 MΩ(2W) | 18998410 | 18998107.94 | 43 | 18987094.742 | 19009725.258 | -15.8992 | 552.5897 | -2.7151 | 0.42149 | +3.5000E+01,PPM,1
10MΩ(2W) | 9999029 | 9998826.214 | 33 | 9998149.134 | 9999908.866 | -20.2806 | 54.9951 | -23.0474 | 0.2099 | +2.7000E+01,PPM,1
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | 24h Spec ohm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 4.8444e-05 | 5e-05 | -5e-05 | 5e-05 | N/A | 8e-06 | PASS | 9.649e-06 | +4.0000E-05,OHM,1
0 Ω | 100 Ω | 0.00051132 | 0.00055 | -0.00055 | 0.00055 | N/A | 2.2e-06 | PASS | 0.00010264 | +4.0000E-05,OHM,1
0 Ω | 1 KΩ | 0.00059651 | 0.0055 | -0.0055 | 0.0055 | N/A | 2.2e-06 | PASS | 0.00012519 | +4.0000E-05,OHM,1
0 Ω | 10 KΩ | 0.0060384 | 5.05 | -5.05 | 5.05 | N/A | 2.2e-06 | PASS | 0.00063664 | +4.0000E-05,OHM,1
0 Ω | 100 KΩ | 0.027307 | 0.55 | -0.55 | 0.55 | N/A | 2.2e-06 | PASS | 0.0022108 | +4.0000E-05,OHM,1
0 Ω | 1 MΩ | 0.21561 | 5.5 | -5.5 | 5.5 | N/A | 2.2e-06 | PASS | 0.032012 | +4.0000E-05,OHM,1
0 Ω | 10 MΩ | 1.7248 | 55 | -55 | 55 | N/A | 2.2e-06 | PASS | 0.32808 | +4.0000E-05,OHM,1
0 Ω | 100 MΩ | 12.7922 | 550 | -550 | 550 | N/A | 2.2e-06 | PASS | 0.1695 | +4.0000E-05,OHM,1
0 Ω | 1 GΩ | 3.6652 | 5500 | -5500 | 5500 | N/A | 2.2e-06 | PASS | 0.18585 | +4.0000E-05,OHM,1
+4.0000E-05,OHM,1
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | 24h Spec ohm | Test Result % of SPEC | MEAS SDEV ohm | +4.0000E-05,OHM,1
0 Ω | 10 Ω | 0.24111 | 0.3 | -0.3 | 0.3 | N/A | 8e-06 | PASS | 9.4864e-05 | +4.0000E-05,OHM,1
0 Ω | 100 Ω | 0.24136 | 0.35 | -0.35 | 0.35 | N/A | 2.2e-06 | PASS | 0.00014504 | +4.0000E-05,OHM,1
0 Ω | 1 KΩ | 0.24125 | 0.4 | -0.4 | 0.4 | N/A | 2.2e-06 | PASS | 7.9285e-05 | +4.0000E-05,OHM,1
0 Ω | 10 KΩ | 0.24901 | 0.4 | -0.4 | 0.4 | N/A | 2.2e-06 | PASS | 0.00015567 | +4.0000E-05,OHM,1
0 Ω | 100 KΩ | 0.27055 | 0.55 | -0.55 | 0.55 | N/A | 2.2e-06 | PASS | 0.0014705 | +4.0000E-05,OHM,1
0 Ω | 1 MΩ | 0.56777 | 5.5 | -5.5 | 5.5 | N/A | 2.2e-06 | PASS | 0.014274 | +4.0000E-05,OHM,1
0 Ω | 10 MΩ | 3.1262 | 55 | -55 | 55 | N/A | 2.2e-06 | PASS | 0.10442 | +4.0000E-05,OHM,1
0 Ω | 100 MΩ | 2.9824 | 550 | -550 | 550 | N/A | 2.2e-06 | PASS | 0.13657 | +5.0000E-05,OHM,365
0 Ω | 1 GΩ | 3.0902 | 5500 | -5500 | 5500 | N/A | 2.2e-06 | PASS | 0.13564 | +5.0000E-05,OHM,365
Note: 1ohm-10k: OCOMP ON, DELAY 0.2s, waiting 10s. Range above 10k OCOMP OFF DELAY default, waiting 100s.
ohm zero: OCOMP OFF DELAY default, waiting 10s
ACV PERFORMANCE TEST
ACV ANA Tset | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured % | 24h Spec % | Test Result % of SPEC | MEAS SDEV ppm
1VAC/50kHz | 1 VAC | 1.0001 | 92 | 0.99401 | 1.0188 | 0.011276 | 0.59 | 1.8819 | 1.0317 | +7.5000E+01,PPM,1
1VAC/1MHz | 1 VAC | 1.0126 | 2000 | 0.924 | 1.1656 | 1.2584 | 7.4 | 16.5573 | 5.8753 | +1.6000E-01,PCT,1
10VAC/40Hz | 10 VAC | 10.0007 | 115 | 9.9529 | 10.1375 | 0.0066224 | 0.47 | 1.3754 | 1.2841 | +4.2000E+01,PPM,1
10VAC/200Hz | 10 VAC | 10.0007 | 52 | 9.9569 | 10.1579 | 0.0068708 | 0.43 | 1.5788 | 1.4632 | +4.2000E+01,PPM,1
10VAC/500Hz | 10 VAC | 10.0006 | 52 | 9.9569 | 10.1376 | 0.0059875 | 0.43 | 1.3758 | 0.7944 | +4.2000E+01,PPM,1
10VAC/50kHz | 10 VAC | 10.0002 | 92 | 9.9409 | 10.0397 | 0.0023788 | 0.59 | 0.397 | 0.9846 | +7.5000E+01,PPM,1
10VAC/1MHz | 10 VAC | 10.1056 | 1900 | 9.2581 | 11.3911 | 1.0558 | 7.4 | 13.9106 | 1.004 | +1.5200E-01,PCT,1
ACV SYNC Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured % | 24h Spec % | Test Result % of SPEC | MEAS SDEV ppm
0.01VAC+DC/10 Hz |  | 0.010001 | 750 | 0.0099865 | 0.010014 | 0.0052149 | 0.06 | 3.8629 | 123.0411 | +6.0000E+02,PPM,1
0.01VAC+DC/20 Hz |  | 0.0099988 | 600 | 0.009988 | 0.010012 | -0.012203 | 0.06 | -10.1696 | 85.9314 | +4.8000E+02,PPM,1
0.01VAC+DC/40 Hz |  | 0.0099991 | 585 | 0.009991 | 0.010009 | -0.0094216 | 0.031 | -10.527 | 28.1934 | +4.7000E+02,PPM,1
0.01VAC+DC/100 Hz |  | 0.0099991 | 585 | 0.009991 | 0.010009 | -0.0092605 | 0.031 | -10.347 | 68.6852 | +4.7000E+02,PPM,1
0.01VAC+DC/1.0 kHz |  | 0.009999 | 585 | 0.009991 | 0.010009 | -0.0099923 | 0.031 | -11.1646 | 65.2699 | +4.7000E+02,PPM,1
0.01VAC+DC/10 kHz |  | 0.01 | 585 | 0.009991 | 0.010009 | 0.0029659 | 0.031 | 3.3139 | 31.9922 | +4.7000E+02,PPM,1
0.01VAC+DC/20 kHz |  | 0.0099992 | 585 | 0.009991 | 0.010009 | -0.0077158 | 0.031 | -8.621 | 26.7038 | +4.7000E+02,PPM,1
0.01VAC+DC/50 kHz |  | 0.0099974 | 720 | 0.0099817 | 0.010018 | -0.026239 | 0.111 | -14.3384 | 77.6609 | +5.7000E+02,PPM,1
0.01VAC+DC/100 kHz |  | 0.0099801 | 1100 | 0.0099379 | 0.010062 | -0.19897 | 0.511 | -32.0403 | 80.3017 | +9.0000E+02,PPM,1
0.01VAC+DC/300 kHz |  | 0.0098416 | 2200 | 0.009576 | 0.010424 | -1.584 | 4.02 | -37.3575 | 49.7323 | +1.3000E-01,PCT,1
0.01VAC+DC/500 kHz |  | 0.0096326 | 3900 | 0.006756 | 0.013244 | -3.6741 | 32.05 | -11.3257 | 43.7804 | +3.1000E-01,PCT,1
0.01VAC+DC/1 MHz |  | 0.0087636 | 5400 | 0.006741 | 0.013259 | -12.3639 | 32.05 | -37.9376 | 66.6654 | +4.4000E-01,PCT,1
0.03VAC+DC/10 Hz |  | 0.030023 | 750 | 0.029971 | 0.030029 | 0.07518 | 0.020333 | 78.8606 | 172.0052 | +6.0000E+02,PPM,1
0.03VAC+DC/20 Hz |  | 0.030015 | 750 | 0.029971 | 0.030029 | 0.05127 | 0.020333 | 53.7794 | 125.5542 | +3.1333E+02,PPM,1
0.03VAC+DC/40 Hz |  | 0.030022 | 351.6667 | 0.029985 | 0.030015 | 0.07442 | 0.013667 | 152.3967 | 238.1168 | +3.0333E+02,PPM,1
0.03VAC+DC/100 Hz |  | 0.030019 | 351.6667 | 0.029985 | 0.030015 | 0.062796 | 0.013667 | 128.5922 | 192.667 | +3.0333E+02,PPM,1
0.03VAC+DC/1.0 kHz |  | 0.030019 | 351.6667 | 0.029985 | 0.030015 | 0.062928 | 0.013667 | 128.8632 | 211.0693 | +3.0333E+02,PPM,1
0.03VAC+DC/10 kHz |  | 0.03002 | 351.6667 | 0.029983 | 0.030017 | 0.065132 | 0.020667 | 116.654 | 171.3815 | +3.0333E+02,PPM,1
0.03VAC+DC/20 kHz |  | 0.030018 | 351.6667 | 0.029983 | 0.030017 | 0.059893 | 0.020667 | 107.2706 | 124.5434 | +3.0333E+02,PPM,1
0.03VAC+DC/50 kHz |  | 0.030016 | 386.6667 | 0.029977 | 0.030023 | 0.053164 | 0.036667 | 70.5723 | 150.6083 | +4.0333E+02,PPM,1
0.03VAC+DC/100 kHz |  | 0.030047 | 1166.6667 | 0.029939 | 0.030061 | 0.1578 | 0.086667 | 77.6057 | 298.7833 | +9.6667E+02,PPM,1
0.03VAC+DC/300 kHz |  | 0.029953 | 1683.3333 | 0.02985 | 0.03015 | -0.1553 | 0.33333 | -30.9576 | 249.3759 | +1.3667E-01,PCT,1
0.03VAC+DC/500 kHz |  | 0.029912 | 1140 | 0.029656 | 0.030344 | -0.29488 | 1.0333 | -25.7015 | 243.5388 | +1.9333E-01,PCT,1
0.03VAC+DC/1 MHz |  | 0.029855 | 4700 | 0.029549 | 0.030451 | -0.48227 | 1.0333 | -32.0799 | 172.0714 | +3.9000E-01,PCT,1
0.10VAC+DC/10 Hz |  | 0.10001 | 400 | 0.099949 | 0.10005 | 0.0072493 | 0.011 | 14.2143 | 41.6876 | +3.2000E+02,PPM,1
0.10VAC+DC/20 Hz |  | 0.099999 | 180 | 0.099971 | 0.10003 | -0.00050866 | 0.011 | -1.754 | 63.9354 | +1.5000E+02,PPM,1
0.10VAC+DC/40 Hz |  | 0.099999 | 165 | 0.099975 | 0.10003 | -0.0012486 | 0.009 | -4.8964 | 62.6481 | +1.4000E+02,PPM,1
0.10VAC+DC/100 Hz |  | 0.099999 | 165 | 0.099975 | 0.10003 | -0.0009341 | 0.009 | -3.6631 | 62.2819 | +1.4000E+02,PPM,1
0.10VAC+DC/1.0 kHz |  | 0.099999 | 165 | 0.099975 | 0.10003 | -0.0011272 | 0.009 | -4.4202 | 43.3995 | +1.4000E+02,PPM,1
0.10VAC+DC/10 kHz |  | 0.1 | 165 | 0.099968 | 0.10003 | 0.0042851 | 0.016 | 13.1849 | 31.4927 | +1.4000E+02,PPM,1
0.10VAC+DC/20 kHz |  | 0.1 | 165 | 0.099968 | 0.10003 | -0.00020584 | 0.016 | -0.63335 | 33.6821 | +1.4000E+02,PPM,1
0.10VAC+DC/50 kHz |  | 0.099996 | 300 | 0.099938 | 0.10006 | -0.0042433 | 0.032 | -6.844 | 27.5986 | +2.4000E+02,PPM,1
0.10VAC+DC/100 kHz |  | 0.099969 | 700 | 0.099848 | 0.10015 | -0.030604 | 0.082 | -20.1344 | 62.3587 | +5.7000E+02,PPM,1
0.10VAC+DC/300 kHz |  | 0.099775 | 1150 | 0.099575 | 0.10043 | -0.22545 | 0.31 | -53.0476 | 50.4604 | +9.0000E+02,PPM,1
0.10VAC+DC/500 kHz |  | 0.099622 | 440 | 0.098946 | 0.10105 | -0.37771 | 1.01 | -35.8359 | 49.7508 | +1.3500E-01,PCT,1
0.10VAC+DC/1 MHz |  | 0.099513 | 3300 | 0.09866 | 0.10134 | -0.48677 | 1.01 | -36.3259 | 38.9654 | +2.8500E-01,PCT,1
0.30VAC+DC/10 Hz |  | 0.30003 | 416.6667 | 0.29981 | 0.30019 | 0.0085028 | 0.020333 | 13.7142 | 23.4024 | +3.3333E+02,PPM,1
0.30VAC+DC/20 Hz |  | 0.30002 | 161.6667 | 0.29989 | 0.30011 | 0.0071336 | 0.020333 | 19.5441 | 34.9079 | +1.2500E+02,PPM,1
0.30VAC+DC/40 Hz |  | 0.30001 | 78.3333 | 0.29994 | 0.30006 | 0.0042424 | 0.013667 | 19.7322 | 24.6709 | +6.3667E+01,PPM,1
0.30VAC+DC/100 Hz |  | 0.30001 | 78.3333 | 0.29994 | 0.30006 | 0.0026256 | 0.013667 | 12.2122 | 28.8972 | +6.3667E+01,PPM,1
0.30VAC+DC/1.0 kHz |  | 0.30002 | 78.3333 | 0.29994 | 0.30006 | 0.0062188 | 0.013667 | 28.9247 | 20.9742 | +6.3667E+01,PPM,1
0.30VAC+DC/10 kHz |  | 0.3 | 78.3333 | 0.29991 | 0.30009 | -0.00067737 | 0.020667 | -2.3767 | 19.7377 | +6.3667E+01,PPM,1
0.30VAC+DC/20 kHz |  | 0.29998 | 78.3333 | 0.29991 | 0.30009 | -0.006971 | 0.020667 | -24.4595 | 12.4315 | +6.3667E+01,PPM,1
0.30VAC+DC/50 kHz |  | 0.30001 | 120 | 0.29985 | 0.30015 | 0.0035389 | 0.036667 | 7.2718 | 20.6153 | +9.8333E+01,PPM,1
0.30VAC+DC/100 kHz |  | 0.30004 | 253.3333 | 0.29966 | 0.30034 | 0.01444 | 0.086667 | 12.8928 | 35.9287 | +2.0000E+02,PPM,1
0.30VAC+DC/300 kHz |  | 0.30033 | 713.3333 | 0.29879 | 0.30121 | 0.10879 | 0.33333 | 26.8839 | 18.3382 | +5.6667E+02,PPM,1
0.30VAC+DC/500 kHz |  | 0.30085 | 1833.3333 | 0.29635 | 0.30365 | 0.28454 | 1.0333 | 23.3871 | 26.9204 | +1.4667E-01,PCT,1
0.30VAC+DC/1 MHz |  | 0.30222 | 2933.3333 | 0.29602 | 0.30398 | 0.73988 | 1.0333 | 55.7701 | 15.8169 | +2.3000E-01,PCT,1
1.00VAC+DC/10 Hz |  | 1.0001 | 300 | 0.99959 | 1.0004 | 0.0064739 | 0.011 | 15.79 | 12.9444 | +2.4000E+02,PPM,1
1.00VAC+DC/20 Hz |  | 1 | 115 | 0.99977 | 1.0002 | 0.0027114 | 0.011 | 12.0507 | 9.3672 | +9.0000E+01,PPM,1
1.00VAC+DC/40 Hz |  | 1 | 55 | 0.99985 | 1.0001 | 0.0028339 | 0.009 | 19.5441 | 3.358 | +4.5000E+01,PPM,1
1.00VAC+DC/100 Hz |  | 1 | 55 | 0.99985 | 1.0001 | 0.0012878 | 0.009 | 8.8814 | 7.2376 | +4.5000E+01,PPM,1
1.00VAC+DC/1.0 kHz |  | 1 | 55 | 0.99985 | 1.0001 | 0.0031214 | 0.009 | 21.5269 | 7.4643 | +4.5000E+01,PPM,1
1.00VAC+DC/10 kHz |  | 0.99996 | 55 | 0.99978 | 1.0002 | -0.0037637 | 0.016 | -17.5058 | 5.1533 | +4.5000E+01,PPM,1
1.00VAC+DC/20 kHz |  | 0.99992 | 55 | 0.99978 | 1.0002 | -0.0075218 | 0.016 | -34.985 | 5.0624 | +4.5000E+01,PPM,1
1.00VAC+DC/50 kHz |  | 0.99999 | 92 | 0.99959 | 1.0004 | -0.00096482 | 0.032 | -2.3418 | 4.6643 | +7.5000E+01,PPM,1
1.00VAC+DC/100 kHz |  | 1.0001 | 124 | 0.99906 | 1.0009 | 0.0051679 | 0.082 | 5.4745 | 9.2007 | +1.3000E+02,PPM,1
1.00VAC+DC/300 kHz |  | 1.0011 | 480 | 0.99642 | 1.0036 | 0.10795 | 0.31 | 30.154 | 7.7784 | +3.8000E+02,PPM,1
1.00VAC+DC/500 kHz |  | 1.0028 | 1250 | 0.98865 | 1.0113 | 0.28305 | 1.01 | 24.938 | 8.2712 | +1.0000E+03,PPM,1
1.00VAC+DC/1 MHz |  | 1.0076 | 2000 | 0.9879 | 1.0121 | 0.76367 | 1.01 | 63.1129 | 7.9338 | +1.6000E-01,PCT,1
3.00VAC+DC/10 Hz |  | 3.0003 | 416.6667 | 2.9981 | 3.0019 | 0.0088321 | 0.020333 | 14.2454 | 18.458 | +3.3333E+02,PPM,1
3.00VAC+DC/20 Hz |  | 3.0002 | 161.6667 | 2.9989 | 3.0011 | 0.0067068 | 0.020333 | 18.3749 | 16.2539 | +1.2500E+02,PPM,1
3.00VAC+DC/40 Hz |  | 3.0001 | 68.3333 | 2.9994 | 3.0006 | 0.004503 | 0.013667 | 21.9657 | 6.2037 | +5.3667E+01,PPM,1
3.00VAC+DC/100 Hz |  | 3.0001 | 68.3333 | 2.9994 | 3.0006 | 0.0035901 | 0.013667 | 17.5125 | 14.3302 | +5.3667E+01,PPM,1
3.00VAC+DC/1.0 kHz |  | 3.0002 | 68.3333 | 2.9994 | 3.0006 | 0.0053093 | 0.013667 | 25.899 | 7.8372 | +5.3667E+01,PPM,1
3.00VAC+DC/10 kHz |  | 3 | 68.3333 | 2.9992 | 3.0008 | -0.000833 | 0.020667 | -3.0291 | 4.7553 | +5.3667E+01,PPM,1
3.00VAC+DC/20 kHz |  | 2.9999 | 68.3333 | 2.9992 | 3.0008 | -0.0027514 | 0.020667 | -10.0052 | 4.6647 | +5.3667E+01,PPM,1
3.00VAC+DC/50 kHz |  | 2.9999 | 120 | 2.9985 | 3.0015 | -0.002526 | 0.036667 | -5.1904 | 16.0067 | +9.8333E+01,PPM,1
3.00VAC+DC/100 kHz |  | 2.9992 | 193.3333 | 2.9968 | 3.0032 | -0.02715 | 0.086667 | -25.6134 | 22.0655 | +1.5667E+02,PPM,1
3.00VAC+DC/300 kHz |  | 2.9961 | 566.6667 | 2.9883 | 3.0117 | -0.13021 | 0.33333 | -33.3859 | 8.6569 | +4.5000E+02,PPM,1
3.00VAC+DC/500 kHz |  | 2.9993 | 1833.3333 | 2.9635 | 3.0365 | -0.022155 | 1.0333 | -1.821 | 9.2993 | +1.4667E-01,PCT,1
3.00VAC+DC/1 MHz |  | 3.0187 | 2833.3333 | 2.9605 | 3.0395 | 0.62409 | 1.0333 | 47.3991 | 9.0822 | +2.2667E-01,PCT,1
10.0VAC+DC/10 Hz |  | 10.0006 | 300 | 9.9959 | 10.0041 | 0.0057875 | 0.011 | 14.1159 | 15.1915 | +2.4000E+02,PPM,1
10.0VAC+DC/20 Hz |  | 10.0003 | 300 | 9.9959 | 10.0041 | 0.002502 | 0.011 | 6.1024 | 9.4142 | +9.0000E+01,PPM,1
10.0VAC+DC/40 Hz |  | 10.0002 | 52 | 9.9986 | 10.0014 | 0.0024895 | 0.009 | 17.5317 | 2.1232 | +4.2000E+01,PPM,1
10.0VAC+DC/100 Hz |  | 10.0001 | 52 | 9.9986 | 10.0014 | 0.0011809 | 0.009 | 8.3162 | 2.8693 | +4.2000E+01,PPM,1
10.0VAC+DC/1.0 kHz |  | 10.0003 | 52 | 9.9986 | 10.0014 | 0.0028668 | 0.009 | 20.1887 | 2.9435 | +4.2000E+01,PPM,1
10.0VAC+DC/10 kHz |  | 9.9997 | 52 | 9.9979 | 10.0021 | -0.0032777 | 0.016 | -15.4609 | 1.9708 | +4.2000E+01,PPM,1
10.0VAC+DC/20 kHz |  | 9.9995 | 52 | 9.9979 | 10.0021 | -0.0046994 | 0.016 | -22.167 | 2.1067 | +4.2000E+01,PPM,1
10.0VAC+DC/50 kHz |  | 9.9993 | 92 | 9.9959 | 10.0041 | -0.0067794 | 0.032 | -16.4549 | 4.0179 | +7.5000E+01,PPM,1
10.0VAC+DC/100 kHz |  | 9.9965 | 135 | 9.9904 | 10.0096 | -0.034585 | 0.082 | -36.2144 | 7.2177 | +1.1000E+02,PPM,1
10.0VAC+DC/300 kHz |  | 9.9869 | 380 | 9.9652 | 10.0348 | -0.13131 | 0.31 | -37.7341 | 5.9141 | +3.1000E+02,PPM,1
10.0VAC+DC/500 kHz |  | 9.9977 | 1250 | 9.8865 | 10.1135 | -0.022858 | 1.01 | -2.0139 | 5.5342 | +1.0000E+03,PPM,1
10.0VAC+DC/1 MHz |  | 10.0641 | 1900 | 9.88 | 10.12 | 0.64083 | 1.01 | 53.4025 | 5.0237 | +1.5200E-01,PCT,1
30.0VAC+DC/10 Hz |  | 30.0014 | 250.1667 | 29.9825 | 30.0175 | 0.0047236 | 0.033333 | 8.0953 | 29.7496 | +3.3333E+02,PPM,1
30.0VAC+DC/20 Hz |  | 30.0003 | 95.0667 | 29.9871 | 30.0129 | 0.0010506 | 0.033333 | 2.4523 | 16.9814 | +1.2500E+02,PPM,1
30.0VAC+DC/40 Hz |  | 30.0003 | 57.0233 | 29.9903 | 30.0097 | 0.0010239 | 0.026667 | 3.1631 | 6.2746 | +6.5000E+01,PPM,1
30.0VAC+DC/100 Hz |  | 30 | 57.0233 | 29.9903 | 30.0097 | 9.6767e-05 | 0.026667 | 0.29895 | 10.601 | +6.5000E+01,PPM,1
30.0VAC+DC/1.0 kHz |  | 30.0003 | 57.0233 | 29.9903 | 30.0097 | 0.0011164 | 0.026667 | 3.4491 | 6.2427 | +6.5000E+01,PPM,1
30.0VAC+DC/10 kHz |  | 29.9995 | 57.0233 | 29.9903 | 30.0097 | -0.0016843 | 0.026667 | -5.2035 | 6.1272 | +6.5000E+01,PPM,1
30.0VAC+DC/20 kHz |  | 29.9987 | 57.0233 | 29.9903 | 30.0097 | -0.0044199 | 0.026667 | -13.6548 | 6.3366 | +6.5000E+01,PPM,1
30.0VAC+DC/50 kHz |  | 29.9996 | 90.04 | 29.9866 | 30.0134 | -0.0014904 | 0.035667 | -3.3364 | 9.6823 | +1.0333E+02,PPM,1
30.0VAC+DC/100 kHz |  | 29.9947 | 160.1 | 29.9572 | 30.0428 | -0.017514 | 0.12667 | -12.2752 | 15.893 | +2.0333E+02,PPM,1
30.0VAC+DC/300 kHz |  | 29.9929 | 900.0667 | 29.843 | 30.157 | -0.023749 | 0.43333 | -4.538 | 5.9992 | +1.2333E-01,PCT,1
30.0VAC+DC/500 kHz |  | 30.0281 | 5001.6667 | 29.3899 | 30.6101 | 0.093619 | 1.5333 | 4.6038 | 27.893 | +5.3333E-01,PCT,1
100.0VAC+DC/10 Hz |  | 100.0011 | 250.05 | 99.951 | 100.049 | 0.0011013 | 0.024 | 2.2473 | 17.1907 | +2.4000E+02,PPM,1
100.0VAC+DC/20 Hz |  | 99.9984 | 95.02 | 99.9665 | 100.0335 | -0.0015602 | 0.024 | -4.657 | 9.4661 | +9.0000E+01,PPM,1
100.0VAC+DC/40 Hz |  | 99.9988 | 57.007 | 99.9723 | 100.0277 | -0.0012364 | 0.022 | -4.4633 | 2.1444 | +5.1000E+01,PPM,1
100.0VAC+DC/100 Hz |  | 99.9975 | 57.007 | 99.9723 | 100.0277 | -0.0024723 | 0.022 | -8.9251 | 1.5545 | +5.1000E+01,PPM,1
100.0VAC+DC/1.0 kHz |  | 99.9991 | 57.007 | 99.9723 | 100.0277 | -0.00088868 | 0.022 | -3.2081 | 1.144 | +5.1000E+01,PPM,1
100.0VAC+DC/10 kHz |  | 99.9976 | 57.007 | 99.9723 | 100.0277 | -0.0023923 | 0.022 | -8.6363 | 2.1719 | +5.1000E+01,PPM,1
100.0VAC+DC/20 kHz |  | 99.9956 | 57.007 | 99.9723 | 100.0277 | -0.0044425 | 0.022 | -16.0375 | 1.9672 | +5.1000E+01,PPM,1
100.0VAC+DC/50 kHz |  | 99.996 | 50.12 | 99.958 | 100.042 | -0.004014 | 0.037 | -9.5543 | 3.7853 | +8.0000E+01,PPM,1
100.0VAC+DC/100 kHz |  | 99.9748 | 160.03 | 99.862 | 100.138 | -0.025236 | 0.122 | -18.2867 | 3.7206 | +1.4500E+02,PPM,1
300.0VAC+DC/100 Hz |  | 300.0026 | 70.0133 | 299.839 | 300.161 | 0.0008758 | 0.046667 | 1.6319 | 6.3286 | +8.8333E+01,PPM,1
300.0VAC+DC/200 Hz |  | 300.0039 | 70.0133 | 299.839 | 300.161 | 0.0012917 | 0.046667 | 2.4068 | 3.8635 | +8.8333E+01,PPM,1
300.0VAC+DC/400 Hz |  | 300.0048 | 70.0133 | 299.839 | 300.161 | 0.0016047 | 0.046667 | 2.99 | 4.944 | +8.8333E+01,PPM,1
300.0VAC+DC/500 Hz |  | 300.0055 | 70.0133 | 299.839 | 300.161 | 0.0018499 | 0.046667 | 3.4469 | 4.9068 | +8.8333E+01,PPM,1
300.0VAC+DC/1 kHz |  | 300.0114 | 70.0133 | 299.839 | 300.161 | 0.0038081 | 0.046667 | 7.0956 | 3.3291 | +8.8333E+01,PPM,1
750.0VAC+DC/1 kHz |  | 750.0386 | 70.0053 | 749.6275 | 750.3725 | 0.0051527 | 0.042667 | 10.3745 | 3.8185 | +8.0333E+01,PPM,1
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | 24h Spec ppm | Test Result % of SPEC | MEAS SDEV ppm
0 µADC | 0 | 1.2825e-13 |  |  |  | N/A |  |  | 1661424.0462 | +6.0000E-09,A,1
50 nADC | 5e-08 | 3.2063e-15 | 140040 | 4.2957e-08 | 5.7042e-08 | -999999.9359 | 810 | -709.9751 | 76678078.8458 | +6.0016E-09,A,1
100 nADC | 1e-07 | -8.9777e-14 | 70040 | 9.2955e-08 | 1.0704e-07 | -1000000.8978 | 410 | -1419.4477 | -2664127.7242 | +6.0032E-09,A,1
-100 nADC | -1e-07 | 9.8113e-13 | 70040 | -9.2955e-08 | -1.0704e-07 | -1000009.8113 | 410 | -1419.4603 | 746230.8959 | +6.0032E-09,A,1
-50 nADC | -5e-08 | 2.3214e-12 | 140040 | -4.2957e-08 | -5.7042e-08 | -1000046.4273 | 810 | -710.0081 | 333492.0367 | +6.0016E-09,A,1
0 µADC | 0 | -4.6869e-12 |  |  |  | N/A |  |  | -396259.1165 | +6.0000E-09,A,1
0.5 µADC | 5e-07 | -4.224e-12 | 14040 | 4.9293e-07 | 5.0706e-07 | -1000008.448 | 90 | -7077.2006 | -188484.0824 | +6.0160E-09,A,1
1 µADC | 1e-06 | -4.3359e-12 | 7040 | 9.9291e-07 | 1.0071e-06 | -1000004.3359 | 50 | -14104.4335 | -311560.8456 | +6.0320E-09,A,1
-1 µADC | -1e-06 | -2.4033e-12 | 7040 | -9.9291e-07 | -1.0071e-06 | -999997.5967 | 50 | -14104.3385 | -378394.3422 | +6.0320E-09,A,1
-0.5 µADC | -5e-07 | -2.5653e-12 | 14040 | -4.9293e-07 | -5.0706e-07 | -999994.8695 | 90 | -7077.1045 | -682847.1549 | +6.0160E-09,A,1
0 µADC | 0 | -1.7869e-11 |  |  |  | N/A |  |  | -234213.0635 | +6.0000E-09,A,1
5 µADC | 5e-06 | -1.6941e-11 | 1440 | 4.9927e-06 | 5.0073e-06 | -1000003.3882 | 24 | -68306.2424 | -277082.1886 | +6.1600E-09,A,1
10 µADC | 1e-05 | -2.0057e-11 | 740 | 9.9924e-06 | 1.0008e-05 | -1000002.0057 | 17 | -132100.6613 | -361144.8369 | +6.3200E-09,A,1
-10 µADC | -1e-05 | -1.2863e-11 | 740 | -9.9924e-06 | -1.0008e-05 | -999998.7137 | 17 | -132100.2264 | -464884.8022 | +6.3200E-09,A,1
-5 µADC | -5e-06 | -1.7173e-11 | 1440 | -4.9927e-06 | -5.0073e-06 | -999996.5654 | 24 | -68305.7763 | -325804.7317 | +6.1600E-09,A,1
0 µADC | 0 | -1.476e-10 |  |  |  | N/A |  |  | -327774.4057 | +6.0000E-09,A,1
50 µADC | 5e-05 | -1.4997e-10 | 180 | 4.999e-05 | 5.001e-05 | -1000002.9995 | 22 | -495050.9898 | -217301.8188 | +1.5200E+02,PPM,1
100 µADC | 0.0001 | -1.7517e-10 | 110 | 9.9987e-05 | 0.00010001 | -1000001.7517 | 16 | -793652.1839 | -249271.905 | +9.2000E+01,PPM,1
-100 µADC | -0.0001 | -1.3548e-10 | 110 | -9.9987e-05 | -0.00010001 | -999998.6452 | 16 | -793649.7184 | -244514.2113 | +9.2000E+01,PPM,1
-50 µADC | -5e-05 | -1.0696e-10 | 180 | -4.999e-05 | -5.001e-05 | -999997.8609 | 22 | -495048.446 | -433869.7221 | +1.5200E+02,PPM,1
0 µADC | 0 | -1.2433e-09 |  |  |  | N/A |  |  | -419145.3327 | +6.0000E-09,A,1
0.5 mADC | 0.0005 | -1.4945e-09 | 46 | 0.00049997 | 0.00050003 | -1000002.9889 | 18 | -1562504.6702 | -219465.647 | +3.9000E+01,PPM,1
1 mADC | 0.001 | -1.2102e-09 | 38 | 0.00099995 | 0.0010001 | -1000001.2102 | 14 | -1923079.2504 | -267268.5172 | +3.2000E+01,PPM,1
-1 mADC | -0.001 | -8.9284e-10 | 38 | -0.00099995 | -0.0010001 | -999999.1072 | 14 | -1923075.2061 | -365738.1209 | +3.2000E+01,PPM,1
-0.5 mADC | -0.0005 | -1.0323e-09 | 46 | -0.00049997 | -0.00050003 | -999997.9353 | 18 | -1562496.7739 | -326523.4213 | +3.9000E+01,PPM,1
0 mADC | 0 | -7.8478e-09 |  |  |  | N/A |  |  | -273661.1871 | +6.0000E-09,A,1
5 mADC | 0.005 | -9.5898e-09 | 40 | 0.0049997 | 0.0050003 | -1000001.918 | 18 | -1724141.2379 | -312167.5352 | +3.3000E+01,PPM,1
10 mADC | 0.01 | -1.0607e-08 | 35 | 0.0099995 | 0.01 | -1000001.0607 | 14 | -2040818.4913 | -150210.8153 | +2.9000E+01,PPM,1
-10 mADC | -0.01 | -1.1177e-08 | 35 | -0.0099995 | -0.01 | -999998.8823 | 14 | -2040814.0456 | -212544.5994 | +2.9000E+01,PPM,1
-5 mADC | -0.005 | -1.0987e-08 | 40 | -0.0049997 | -0.0050003 | -999997.8026 | 18 | -1724134.1425 | -199083.0782 | +3.3000E+01,PPM,1
0 mADC | 0 | -6.2035e-08 |  |  |  | N/A |  |  | -337985.2905 | +6.0000E-09,A,1
50 mADC | 0.05 | -4.4538e-08 | 40.016 | 0.049996 | 0.050004 | -1000000.8908 | 33 | -1369564.0555 | -374905.5059 | +4.9000E+01,PPM,1
100 mADC | 0.1 | -4.6129e-08 | 40.008 | 0.099993 | 0.10001 | -1000000.4613 | 29 | -1449108.0183 | -543308.664 | +4.2000E+01,PPM,1
-100 mADC | -0.1 | -3.9766e-08 | 40.008 | -0.099993 | -0.10001 | -999999.6023 | 29 | -1449106.7736 | -869048.0104 | +4.2000E+01,PPM,1
-50 mADC | -0.05 | -4.6129e-08 | 40.016 | -0.049996 | -0.050004 | -999999.0774 | 33 | -1369561.572 | -660244.5328 | +4.9000E+01,PPM,1
0 ADC | 0 | -1.0316e-06 |  |  |  | N/A |  |  | -189480.937 | +6.0000E-09,A,1
0.5 ADC | 0.5 | -6.7586e-07 | 60.03 | 0.49991 | 0.50009 | -1000001.3517 | 120 | -555463.7292 | -396657.1225 | +7.4000E+01,PPM,1
1.0 ADC | 1 | -9.7822e-07 | 60.015 | 0.99983 | 1.0002 | -1000000.9782 | 110 | -588183.971 | -306465.379 | +6.2000E+01,PPM,1
-1.0 ADC | -1 | -7.8258e-07 | 60.015 | -0.99983 | -1.0002 | -999999.2174 | 110 | -588182.9353 | -379809.229 | +6.2000E+01,PPM,1
-0.5 ADC | -0.5 | -6.0472e-07 | 60.03 | -0.49991 | -0.50009 | -999998.7906 | 120 | -555462.3066 | -526544.5234 | +7.4000E+01,PPM,1
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured % | 24h Spec % | Test Result % of SPEC | MEAS SDEV ppm
10.0 µA/50Hz | 1e-05 | 5.751e-08 | 1120 | 9.9528e-06 | 1.0047e-05 | -99.4249 | 0.36 | -21064.5971 | 20090.3975 | +9.0000E-09,A,1
10.0 µA/60Hz | 1e-05 | 5.9422e-08 | 1120 | 9.9528e-06 | 1.0047e-05 | -99.4058 | 0.36 | -21060.5464 | 28734.2685 | +9.0000E-09,A,1
10.0 µA/1kHz | 1e-05 | 5.9774e-08 | 1120 | 9.9528e-06 | 1.0047e-05 | -99.4023 | 0.36 | -21059.8006 | 23671.2099 | +9.0000E-09,A,1
100 µA/50Hz | 0.0001 | 5.6189e-08 | 220 | 9.9888e-05 | 0.00010011 | -99.9438 | 0.09 | -89235.5455 | 21904.6901 | +1.8000E+02,PPM,1
100 µA/60Hz | 0.0001 | 6.0382e-08 | 220 | 9.9888e-05 | 0.00010011 | -99.9396 | 0.09 | -89231.802 | 30545.2991 | +1.8000E+02,PPM,1
100 µA/1kHz | 0.0001 | 5.9729e-08 | 220 | 9.9888e-05 | 0.00010011 | -99.9403 | 0.09 | -89232.3845 | 37958.3124 | +1.8000E+02,PPM,1
1.00 mA/50Hz | 0.001 | 3.8681e-07 | 160 | 0.00099904 | 0.001001 | -99.9613 | 0.08 | -104126.3739 | 31476.9705 | +1.3500E+02,PPM,1
1.00 mA/60Hz | 0.001 | 4.058e-07 | 160 | 0.00099904 | 0.001001 | -99.9594 | 0.08 | -104124.3958 | 33919.2505 | +1.3500E+02,PPM,1
1.00 mA/1kHz | 0.001 | 3.9539e-07 | 160 | 0.00099934 | 0.0010007 | -99.9605 | 0.05 | -151455.244 | 20232.299 | +1.3500E+02,PPM,1
10.0 mA/50Hz | 0.01 | 3.5414e-05 | 160 | 0.0099904 | 0.01001 | -99.6459 | 0.08 | -103797.7671 | 32036.7262 | +1.3500E+02,PPM,1
10.0 mA/60Hz | 0.01 | 3.6947e-05 | 160 | 0.0099904 | 0.01001 | -99.6305 | 0.08 | -103781.7983 | 24308.3758 | +1.3500E+02,PPM,1
10.0 mA/1kHz | 0.01 | 3.7186e-05 | 160 | 0.0099934 | 0.010007 | -99.6281 | 0.05 | -150951.7287 | 42901.1793 | +1.3500E+02,PPM,1
100 mA/50Hz | 0.1 | 3.6321e-05 | 120.03 | 0.099908 | 0.10009 | -99.9637 | 0.08 | -108652.6303 | 41213.8047 | +1.2500E+02,PPM,1
100 mA/60Hz | 0.1 | 3.6965e-05 | 120.03 | 0.099908 | 0.10009 | -99.963 | 0.08 | -108651.9303 | 23849.6487 | +1.2500E+02,PPM,1
100 mA/1kHz | 0.1 | 3.7015e-05 | 120.03 | 0.099938 | 0.10006 | -99.963 | 0.05 | -161222.8199 | 25656.2786 | +1.2500E+02,PPM,1
1.00 A/50Hz | 1 | 0.00037106 | 290.04 | 0.99871 | 1.0013 | -99.9629 | 0.1 | -77488.2127 | 49431.6926 | +2.6500E+02,PPM,1
1.00 A/60Hz | 1 | 0.00035995 | 290.04 | 0.99871 | 1.0013 | -99.964 | 0.1 | -77489.0736 | 41399.8815 | +2.6500E+02,PPM,1
1.00 A/1kHz | 1 | 0.00034924 | 290.04 | 0.99851 | 1.0015 | -99.9651 | 0.12 | -67088.8542 | 38080.5925 | +2.6500E+02,PPM,1
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.03 | 29.56 | 48.97 | 76.19 | 30.2 | 9.09 |  | 194.04
delta Temp degree | 0 | -0.1 | 0.2 | 0.1 | -0.1 | 0 |  | 0.1